Competitive Enzyme Immunoassay Kit for

Quantitative Analysis of sulfamethazine (SM2)

1. Background

Sulfamethazine ( SM2) is a broadly applied antibacterial, which once played important role in controlling and curing animal diseases. These drugs have very serious side effects and will lead to SAs resistance of some bacilus if they exist in human body for a long period. They also have potential carcinogenecity.Doc.235 of Ministry of Agriculture prescribes that the residue limit of SAs is 100(g/kg.

This kit is a new generation product for drug residue detection based on ELISA technology. It is fast, simple, accurate and sensitive. And it requires only 1.5 hours in one operation, which considerably minimizes work intensity and operation error. This kit is documented in Doc.558 of MOA in Oct.2005.

2. Test Principle

This kit is based on indirect-competitive ELISA technology. The microtiter wells are coated with coupling antigen. Sulfamethazin residue in the sample competes with the antigen coated on the microtiter plate for the antibody. After the addition of enzyme labeled anti-antibody, TMB substrate is used to show the color. Absorbance of the sample is negatively related to the SM2 residing in it, after comparing with the Standard Curve, multiplied by the dilution multiple, SM2 residue quantity in the sample can be calculated.

3. Applications

This kit can be used for quantitative and qualitative analysis of SM2 residue in animal tissues (muscle , liver ,fish and shrimp), egg, honey, milk and serum .

4. Cross-reactions

Sulfamethazin (SM2)…………………………… 100.0%

Sulfameter  (SMD)………………………… 10.0%

Sulfamerazine  (SM1)……………………………  5.0%

Sulfamonomethoxine  (SMM)……………………… 0.8%

sulfaquinoxaline (SQX)……………………………… 1.1%

sulfamethoxypyridazine (SMP)……………………… 1.4%

sulfamethoxazole (SMZ)……………………………… 0.1%

sulfadiazine (SD)…………………………………   0.6%

sulfadimethoxine (SDM)………………………0.1%

5. Materials Required

5.1 Equipments:

┅┅Microtiter plate spectrophotometer (450nm/630nm)

┅┅Rotary evaporator or nitrogen drying instruments

┅┅Homogenizer or stamocher

┅┅Shaker

┅┅Gyroscope  or vortex mixer

┅┅Centrifuge

┅┅Analytical balance (inductance: 0.01g)

┅┅Graduated pipette: 10ml

┅┅Rubber pipette bulb

┅┅volumetric flask： 1L

┅┅glass test rube：10ml

┅┅Polystyrene Centrifuge tube: 1.5ml、50ml

┅┅Micropipettes: 20(l～200(l、100(l～1000(l

250(l -multiple

5.2 Reagents:

┅┅Acetonitrile (AR) (For animal tissues)

---- Ethyl acetate (AR)

----N-hexane (AR)

----disodium hydrogen phosphate dodecahydate (AR)

---- potassium dihydrogen phosphate dihydrate (AR)

----Sodium chloride (AR)

---- potassium chloride (AR)

----deionized water

6. Kit Components

1、Microtiter plate with 96 wells coated with coupling antigen

2、Standard solutions(6 bottles×1ml/bottle)

0ppb，1ppb，3ppb，9ppb，27ppb，81ppb

3、High concentration standard solutions (1bottle×1ml)1ppm

4、enzyme-labeled secondary antibody solution 7ml… red cap

5、antibody solution 10ml …………………green cap

6、solution  A  7ml …………………white cap

7、solution  B  7ml …………………red cap

8、stop solution  7ml ……………………yellow cap

9、20×Concentrated wash solution  40ml……transparent cap

10、20×concentrated extraction solution 50ml………blue cap

7. Reagents Preparation:

solution 1: 0.02M PBS1 buffer solution

Weigh 5.36g disodium hydrogen phosphate dodecahydate, 0.20g potassium dihydrogen phosphate dihydrate, 8.00g Sodium chloride and 0.20g potassium chloride, dissolve with deionized water and dilute to 1L ;

solution 2: extraction solution

Dilute the 20×concentrated extraction solution with deionized water in the volume ration of 1:19(or depend on requirement), which will be used for sample extraction. This diluted solution can be conserved for 1 month at 4℃.

solution 3: wash solution

Dilute the 20×concentrated wash solution with deionized water in the volume ration of 1:19(or depend on requirement), which will be used to wash the plates. This diluted solution can be conserved for 1 month at 4℃.

8. Sample Preparations

8.1 Notice and precautions for the users before operation:

(a) Please use one-off tips in the process of experiment, and change the tips when absorb different reagent.

(b) Make sure that all experimental instruments are clean.

(c) Samples that are treated with low limit sample treat method can not be conserved, must be used for assay after diluted with extraction solution.

(d) Samples that are treated with other methods can be conserved at 4℃ for 24 hours.

8.2 Samples of high detection limit :

(a) Animal tissues (chicken, chicken liver, pork, pig liver, eggs)

┅┅Homogenize the egg sample with homogenizer or stamocher at low speed to completely mix the egg white and yolk

┅┅Weigh 2.0±0.05g homogenate to a 50ml polystyrene centrifuge tube, add 8ml Acetonitrile, shake immediately with shaker for 20min, and then centrifuge for 5min,at least 3000g.

┅┅Take 2.5ml of the supernatant liquid to a 10ml clean test tube, dry it with 50～60℃ water bath nitrogen flow.

┅┅Add 1ml n-hexane ,whorl 20s to dissolve the dry leftover, and then add 1ml 0.02 M PBS1 buffer solution(see solution 1),whorl for 1min with gyroscope  or vortex mixer, centrifuge for 10min, at least 3000g.

┅┅Remove the supernatant n-hexane phase, take 20(l of the substrate water phase for assay.

(b)Animal tissue (Fish and shrimp)

┅┅Homogenize the sample with homogenizer or stamocher

┅┅Weigh 2.0±0.05g of the homogenate to a 50ml polystyrene centrifuge tube, add 2ml deionized water, whorl it with gyroscope  or vortex mixer to paste, add 8ml acetonitrile with 10ml Graduated pipette, shake immediately with shaker for 20min,and then centrifuge for 5min, at least 3000g.

┅┅Take 2.5ml of the supernate into a 10ml clean test tube, dry it with 50～60℃ water bath nitrogen flow.

┅┅Add 1ml n-hexane to a 50ml polystyrene centrifuge tube, dissolve the dry leftover by whorling for 20s with gyroscope  or vortex mixer, and then add 1ml 0.02 M PBS1 buffer solution(see solution 1),and then whorl it for 1min with gyroscope  or vortex mixer, centrifuge for 10min,at least 3000g.

┅┅Remove the supernatant n-hexane phase, take 20(l of the substrate water phase for assay.

8.3 samples of low detection limit

(a)  Animal tissue ( muscle and liver samples low in fat)

┅┅Homogenize the sample with homogenizer or stamocher;

┅┅Weigh 2.0±0.05g homogenate to a 50ml polystyrene centrifuge tube, add 10ml extraction solution (see solution 2), shake with shaker for 10min, and then incubate in 37℃ thermotank for 30min, centrifuge for 10min, at least 3000g, at 10℃.

┅┅Take 20(l of the supernate for assay.

(b)  Animal tissue ( muscle and liver samples rich in fat)

┅┅Homogenize the sample with homogenizer or stamocher;

┅┅Weigh 2.0±0.05g homogenate to a 50ml polystyrene centrifuge tube, add 10ml extraction solution (see solution 2) and 5ml n-hexane, shake with shaker for 10min, centrifuge for 15min, at least 3000g, at 10℃.

-----Remove the supernatant n-hexane phase, take 100(l of the substrate phase into a 1.5ml polystyrene centrifuge tube, add 0.02 M PBS1 buffer solution (see solution 1) 100(l, mix completely;

┅┅Take 20(l of the prepared solution for assay.

8.4 Serum samples

┅┅Keep the blood sample at room temperature for 30min, and then centrifuge for 10min, at least 3000g, at 10℃, separate to obtain serum or filter the serum.

┅┅Take 1ml serum to a 10ml clean glass test tube, then add 3ml 0.02 M PBS1 buffer solution (see solution 1), mix adequately.

┅┅Take 20(l of the prepared solution for assay.

8.5 Honey

┅┅Weigh 1.0±0.05g honey sample to a 50ml polystyrene centrifuge tube. Add 2ml 0.02 M PBS1 buffer solution (see solution 1) to dissolve it;

┅┅Add 8ml ethyl acetate , shake for 10min with shaker, centrifuge for 10min, at least 3000g ;

┅┅Take the 4ml of the supernatant liquid into a 10ml clean glass test tube, dry it with 50～60℃ water bath nitrogen flow, then add 0.5ml 0.02 M PBS1 buffer solution (see solution 1) to dissolve the dry leftover.

┅┅Take 20(l of the prepared solution for assay ;

8.6 Urine

┅┅Centrifuge the urine sample for 5min, at least 3000g, at room temperature (20-25℃) until it turns transparent.

┅┅Transfer 1ml transparent urine to a 10ml clean glass test tube, add 3ml extraction solution and mix completely.

┅┅Take 20(l of the prepared solution for assay.

8.7 Milk

┅┅Centrifuge the milk sample for 10min, at least 3000g, at 10℃, then remove the supernatant fat.

┅┅Take 1ml centrifuged milk sample, dilute with 0.02M PBS1 buffer solution in the volume ratio of 1:19 (20(l milk sample + 380(l 0.02 PBS1 buffer solution).

┅┅Take 20(l of the prepared solution for assay.

9. Assay process

9.1 Notice before assay:

9.1.1 Make sure all reagents and micro wells are all at room temperature (20-25℃).

9.1.2 Return all the rest reagents to 2～8℃ immediately after used.

9.1.3 Washing the microwells correctly is an important step in the process of assay; it is the vital factor to the repetitiveness of the ELISA analysis.

9.1.4 Avoid the light and cover the microwells during incubation.

9.2 Assay Steps:

9.2.1 Take all reagents out at room temperature (20-25℃) for more than 30min, homogenize before use.

9.2.2 Get the microwells needed out and return the rest into the zip-lock bag at 2-8℃ immediately.

9.2.3 The diluted wash solution should be rewarmed to be at room temperature before use.

9.2.4 Number: Numbered every microwell positions and all standards and samples should be run in duplicate. Record the standards and samples positions.

9.2.5 Add standard solution/sample: Add 20 µl of standard solution or prepared sample to corresponding wells. Add 50µl of the enzyme labeled secondary antibody solution, then add 80µl of the antibody solution to each well, Mix gently by rocking the plate manually and incubate for 60min at 25℃ with cover.

9.2.6 Wash: Remove the cover gently and pure the liquid out of the wells and rinse the microwells with 250µl diluted wash solution (solution 3) at interval of 10s for 4~5 times. Absorb the residual water with absorbent paper (the rest air bubble can be eliminated with unused tip).

9.2.7 Coloration: Add 50µl solution A and 50µl solution B to each well. Mix gently by rocking the plate manually and incubate for  30min at 25℃ with cover (see 12.8).

9.2.8 Measure: Add 50µl the stop solution to each well. Mix gently by rocking the plate manually and measure the absorbance at 450nm against an air blank (It’s suggested measure with the dual-wavelength of 450/630nm. Read the result within 5min after addition of stop solution. ) (We can also measure by sight without stop solution in short of the ELIASA instrument)

10. Results

There are 2 different methods to determinate the results. Method 1 led to a round estimation, and method 2 lead to definite quantity estimation. (Please notice: the absorption is inversely proportional to the SM2 concentration in the sample.)

10.1 Round estimation

We can get the range of different strengths from the compare of average absorption and standards by sight. For example, the absorption of sample 1 is 0.111, sample 2 is 0.585, and the absorptions of SM2 standard solutions: 1.428 (0ppb); 1.047(1ppb); 0.645(3ppb); 0.337(9ppb); 0.161(27ppb); 0.063(81ppb). So we can say the strength of diluted sample 1 is between 27ppb and 81ppb; and diluted sample 2 between 3ppb to 9ppb. In order to obtain the SM2 actually contained in a sample, the diluted sample results must be further multiplied by the corresponding dilution multiple.

10.2 Definite quantity estimation

(1) The mean values of the absorbance values obtained for the standards and the samples are divided by the absorbance value of the first standard (zero standard ) and multiplied by 100%. The zero standard is thus made equal to 100% and the absorbance values are quoted in percentages.

B

Absorbance (%) = —— ×100%

B0

B ——absorbance standard (or sample)

B0 ——absorbance zero standard

(2) To draw a standard curve: Take the absorbance value of standards as y-axis, semi logarithmic of the concentration of the SM2 standards solution (ppb) as x-axis.

--- The SM2 concentration of each sample (ppb), which can be read from the calibration curve, is multiplied by the corresponding dilution rate of each sample followed, and the actual concentration of sample is obtained.

Please notice:

For evaluation of the ELISA kits special software has been developed for exact and rapid analysis. The analysis software can be ordered on request.

Dilution multiple of samples：

(1) High detection limit samples:

Muscle , liver and egg samples treated with Method (a):2;

Fish and shrimp samples treated with Method(b): 2 ;

(2)Low detection limit samples:

Muscle and liver samples low in fat:  5

Muscle and liver samples rich in fat: 10

(3)Serum samples: 4;

(4) honey samples: 1;

(5) urine samples: 4;

(6) milk samples: 20;

11. Sensitivity, accuracy and precision

Test Sensitivity: 1ppb

Detection limit:

Honey sample……………………………………1ppb

Serum and urine samples…………………………………4ppb

Milk sample……………………………………………20ppb

Detection limit of tissue samples:

High detection limit samples:

Muscle , liver and egg samples……………………………2ppb

Fish and shrimp samples…………………………………2ppb

Low detection limit samples:

Muscle and liver…………………………………………5ppb

Muscle and liver samples rich in fat…………………… 10ppb

Accuracy :

Muscle and liver of chicken and pig…………………75±10%

Egg samples …………………………………………69±10%

honey, milk and serum samples……………………70±10%

Precision:

Variation coefficient of the ELISA kit is less than 10%.

12. Notice

12.1 The mean values of the absorbance values obtained for the standards and the samples will be reduced if the reagents and samples have not been regulated to room temperature (20-25℃).

12.2 Do not allow microwells to dry between steps to avoid unsuccessful repetitiveness and operate the next step immediately after tap the microwells holder.

12.3. Homogenize each reagent before using.

12.4. Keep your skin away from the stop solution for it is the 2M H2SO4 solution.

12.5 Don’t use the kits out of date. Don’t exchange the reagents of different batches, or else it will drop the sensitivity.

12.6 Storage condition:

Keep the ELISA kits at 2-8℃, do not freeze. Seal rest microwell plates Avoid straight sunlight during all incubations.

Covering the microtiter plates is recommended.

12.7 Indications for the reagents going bad:

Substrate solution should be abandoned if it turns colors.

The reagents may be turn bad if the absorbance value (450/630nm) of the zero standard is less than 0.5(A450nm＜0.5).

12.8 The coloration reaction need 20-30min after the addition of solution A and solution B; But you can prolong the incubation time ranges from 35min to more if the color is too light to be determined., never exceed 40min,On the contrary, shorten the incubation time properly.

12.9 The optimal reaction temperature is 25℃. Higher or lower temperature will lead to the changes of sensitivity and absorbance values.

13. Storage condition and storage period

Storage condition: 2-8℃.

Storage period: 6 months