Competitive Enzyme Immunoassay Kit for

Quantitative Analysis of Ractopamine

1. Background

Ractopamine belongs to β-agonists family. β-agonist is a kind of nutrient repartitioning agents, which is a phenylethanolamine ramification structurally and functionally similar to adrenalin and noradrenalin, it can promote animal growth, decrease body fat content and improve lean meat percentage. It has been inhibited to promote animal growth in EU, as well as in China. Because ractopamine is a nonprescription drug, which is much cheaper than clenbuterol, it is being illegally used quite often in production.

This kit is a new product based on ELISA technology, which is fast, easy, accurate and sensitive compared with common instrumental analysis, so it can considerably minimize operation error and work intensity. This kit has been documented in Doc. 558 of Ministry of Agriculture(China, PRC).

2. Test Principle

This kit is based on indirect-competitive ELISA technology. The microtiter wells are coated with coupling antigen. ractopamine residue in the sample competes with the antigen coated on the microtiter plate for the antibody. After the addition of enzyme labeled anti-antibody, TMB substrate is used to show the color. Absorbance of the sample is negatively related to the ractopamine reside in it, after comparing with the Standard Curve, multiplied by the dilution multiple, ractopamine residue quantity in the sample can be calculated.

3. Applications

This kit can be used in quantitative and qualitative analysis of ractopamine residue in urine sample.

4. Cross-reactions

ractopamine…………………………………  100%

dobutamine …………………………………………20%

clenbuterol………………………………………<1%

dopamine hydrochloride………………………………3.4%

salbutamol……………………………………… <1%

5. Materials Required

5.1 Equipments:

┅┅Microtiter plate spectrophotometer (450nm/630nm)

┅┅Rotary Evaporator/Nitrogen drying instrument

┅┅Homogenizer or stamocher

┅┅Shaker

┅┅Gyroscope  or vortex mixer

┅┅Centrifuge

┅┅Analytical balance (inductance: 0.01g)

┅┅Graduated pipette: 10ml

┅┅Rubber pipette bulb

┅┅glass test tube：10ml

┅┅volumetric flask：100ml、500ml

┅┅Polystyrene Centrifuge tube： 50ml

┅┅Micropipettes：20(l～200(l、100(l～1000(l

250ul-multipipette

5.2 Reagents:

---- ethyl acetate (AR)(For urine sample)

----Anhydrous acetonitrile(AR) (For tissue sample)

----n-hexane(AR) (For tissue sample)

----Anhydrous sodium carbonate (AR) (For urine sample)

----sodium bicarbonate (AR) (For urine sample)

---- Disodium hydrogen phosphate dodecahydrate (AR)

---- potassium dihydrogen phosphate dihydrate (AR)

----sodium chloride (AR)

----potassium chloride (AR)

----concentrated hydrochloric acid  (AR)

----deionized water

6. Kit Components

1、Microtiter plate with 96 wells coated with coupling antigen

2、Standard solutions(6 bottles×1ml/bottle)

0ppb，0. 5ppb，1.5ppb，4.5ppb，13.5ppb，40.5ppb

3、high concentration standard solution：(1ml/bottle)  100ppb

4、enzyme-labeled secondary antibody solution  7ml …… red cap

5、antibody work solution  10ml …………green cap

6、solution A 7ml ………………white cap

7、solution B 7ml ……………………red cap

8、stop solution  7ml …………………yellow cap

9、20×concentrated wash solution  40ml…transparent cap

10、20×concentrated extraction solution 50ml …blue cap

7. Reagents Preparation:

Solution 1: pH 10.6 0.1M carbonate buffer solution(for Method I of urine sample)

Weigh 2.56g  anhydrous sodium carbonate and 2.1g sodium bicarbonate, dissolve with deionized water and dilute to 500ml;

Solution 2: extraction solution (for Method II of urine sample)

Dilute the 20×concentrated extraction solution with deionized water in the volume ration of 1:19, which will be used to dilute the urine sample. This diluted extraction solution can be conserved for 1 month at  4℃.

Solution 3: Wash solution

Dilute the 20×concentrated wash solution with deionized water in the volume ration of 1:19, which will be used to rinse the plates This diluted wash solution can be conserved for 1 month at  4℃.

Solution 4: Extraction solution of tissue sample ( for Method IV)

Add 5g sodium chloride into 100ml diluted extraction solution(see solution 2), then add 0.083 ml concentrated  hydrochloric acid to regulate the pH to 2~3, mix completely. (according to the quantity of  samples tested with Method IV )

Solution 5: Extraction solution (phosphate buffer solution) (for Method I and III)

Weigh 10.72g Disodium hydrogen phosphate dodecahydrate, 0.4g potassium dihydrogen phosphate dihydrate, 16g sodium chloride and  0.4g potassium chloride, dissolve with deionized water and dilute to 1L .

8. Sample Preparations

8.1 Notice and precautions for the users before operation:

(a) Please use one-off tips in the process of experiment, and change the tips when absorbing different reagent.

(b) Make sure that all experimental instruments are clean.

(c) For some urine samples have some background interference( the interference value is between stand solution 2 and 3), urine samples treated with extraction method, 2ppb is recommended as the cut off value for positive results.

8.2 Urine sample ( Extraction Method )

┅┅Centrifuge the urine sample at 15℃ for 5min, at least 3000g until the sample is transparent;

----Take 1ml of the supernate into a 10ml glass tube, add 1ml pH 10.6 0.1M carbonate buffer solution, mix and add 4ml   ethyl acetate , shake with shaker, centrifuge the urine sample at 15℃ for 5min, at least 3000g;

----Take 2ml of the supernatant phase into a 10ml clean glass test tube, dry with 50～60℃ nitrogen gas flow;

----Add 1ml extraction solution, whorl for 2min to dissolve the leftover;

----Take 20(l of the prepared solution for assay;

8.3 Urine sample ( Dilution Method )

---- Mix the urine sample with extraction solution in the volume ration of 1:1, take 20(l of the prepared solution for assay;

8.4 Method I for tissue and liver samples

┅┅Homogenize the samples with homogenizer or stamocher;

----Weigh 3.0±0.05g of the homogenate into a 50ml polystyrene centrifuge tube, then add 9ml anhydrous acetonitrile, shake for 10min with shaker, centrifuge at 15℃ for 10min , at least 3000g;

----Transfer 4ml of the supernate into a 10ml clean glass test tube, dry with 50～60℃ nitrogen gas flow;

----Add 1ml n-hexane, whorl for 30s with gyroscope  or vortex mixer;

----Add 1ml extraction solution, whorl for 1min with gyroscope  or vortex mixer, then centrifuge at 15℃ for 5min , at least 3000g, remove the supernatant phase;

Note：

-----Muscle sample: Take 100(l of the substrate phase and mix with 100(l extraction solution;

----Take 20(l of the prepared solution for assay;

-----Liver sample: Take 50(l of the substrate phase and mix with 150(l extraction solution;

----Take 20(l of the prepared solution for assay;

8.5 Method II for tissue and liver samples:

----Homogenize the sample with homogenizer or stamocher;

----Weigh 2.0±0.05g of the homogenate into a 50ml centrifuge tube, add 8ml extraction solution of tissue sample. Whorl for 5min, incubate in 80℃ water bath for 20min, take out and whorl again, then incubate in 80℃ water bath for 20min;

---- centrifuge at 10℃ for 10min , at least 3000g;

----Take 20(l of the prepared solution for assay;

Note: if the supernate is still muddy after centrifuge, then take 1~2ml of it into a centrifuge tube, centrifuge again at high speed and take the transparent supernate for assay;

9. Assay process

9.1 Notice before assay:

9.1.1 Make sure all reagents and microwells are all at room temperature (20-25℃).

9.1.2 Return all the rest reagents to 2～8℃ immediately after used.

9.1.3 Washing the microwells correctly is an important step in the process of assay; it is the vital factor to the reproducibility of the ELISA analysis.

9.1.4 Avoid the light and cover the microwells during incubation.

9.2 Assay Steps:

9.2.1 Take all reagents out at room temperature (20-25℃) for more than 30min, homogenize before use.

9.2.2 Get the microwells needed out and return the rest into the zip-lock bag at 2-8℃ immediately.

9.2.3 The diluted wash solution should be rewarmed to be at room temperature before use.

9.2.4 Number: Numbered every microwell positions and all standards and samples should be run in duplicate. Record the standards and samples positions.

9.2.5 Add standard solution/sample: Add 20 µl of standard solution or prepared sample to corresponding wells. Add 50(l enzyme labeled secondary antibody solution , then add 80µl antibody solution. Mix gently by rocking the plate manually and incubate for 40min at 25℃ with cover.

9.2.6 Wash: Remove the cover gently and pure the liquid out of the wells and rinse the microwells with 250µl diluted wash solution (solution 3) at interval of 10s for 4~5 times. Absorb the residual water with absorbent paper (the rest air bubble can be eliminated with unused tip).

9.2.7  Coloration: Add 50µl solution A and 50µl solution B to each well. Mix gently by rocking the plate manually and incubate for 15min at 25℃ with cover(see 12.8).

9.2.8 Measure: Add 50µl the stop solution to each well. Mix gently by rocking the plate manually and measure the absorbance at 450nm against an air blank (It’s suggested measure with the dual-wavelength of 450/630nm. Read the result within 5min after addition of stop solution. ) (We can also measure by sight without stop solution in short of the ELIASA instrument)

10. Results

There are 2 different methods to determinate the results. Method 1 leads to a round estimation, and method 2 leads to definite quantity estimation. (Please notice: the absorption is inversely proportional to the ractopamine concentration in the sample.)

10.1 Round estimation

We can get the range of different strengths from the compare of average absorption and standards by sight. For example, the absorption of sample 1 is 0.358, sample 2 is 0.758, and the absorptions of ractopamine standard solutions: 1.952 (0ppb); 1.529(0.5ppb); 1.072(1.5ppb); 0.547(4.5ppb); 0.224(13.5ppb); 0.094(40.5ppb). So we can say the strength of diluted sample 1 is between 4.5ppb and 13.5ppb; and diluted sample 2 between 1.5 ppb to 4.5 ppb. In order to obtain the ractopamine actually contained in a sample, the diluted sample results must be further multiplied by the corresponding Dilution factor.

10.2 Definite quantity estimation

(1) The mean values of the absorbance values obtained for the standards and the samples are divided by the absorbance value of the first standard (zero standard ) and multiplied by 100%. The zero standard is thus made equal to 100% and the absorbance values are quoted in percentages.

B

Absorbance (%) = —— ×100%

B0

B ——absorbance standard (or sample)

B0 ——absorbance zero standard

(2) To draw a standard curve: Take the absorbance value of standards as y-axis, semi logarithmic of the concentration of the ractopamine standards solution (ppb) as x-axis.

--- The ractopamine concentration of each sample (ppb), which can be read from the calibration curve, is multiplied by the corresponding Dilution factor of each sample followed, and the actual concentration of sample is obtained.

Please notice:

For evaluation of the ELISA kits special software has been developed for exact and rapid analysis. The analysis software can be ordered on request.

Dilution factor of samples：

Urine samples：2( Extraction method, dilute method )

Muscle samples: 2 (Method III)

Liver samples: 4 (Method III)

Muscle and liver samples : 5 (Method IV)

11. Sensitivity, accuracy and precision

Test Sensitivity:0.5ppb

Detection limit:

Urine samples……………………………………1ppb

Muscle sample………………………………………1ppb

Liver sample…………………………………2ppb

(Method I for tissue and liver samples)

Muscle and liver samples………………………… 2.5 ppb

(Method II for tissue and liver samples)

Accuracy:

Urine sample……………………………………80±10%

Tissue sample(muscle and liver)……………70±10%

Precision:

Variation coefficient of the ELISA kit is less than 10%.

12. Notice

12.1 The mean values of the absorbance values obtained for the standards and the samples will be reduced if the reagents and samples have not been regulated to room temperature ( 20-25℃).

12.2 Do not allow microwells to dry between steps to avoid unsuccessful reproducibility and operate the next step immediately after tap the microwells holder.

12.3. Homogenize each reagent before using.

12.4. Keep your skin away from the stop solution for it is the 2M H2SO4 solution.

12.5 Don’t use the kits out of date. Don’t exchange the reagents of different batches, or else it will drop the sensitivity.

12.6 Storage condition:

Keep the ELISA kits at 2-8℃,do not freeze. Seal rest microwell plates, Avoid straight sunlight during all incubations. Covering the microtiter plates is recommended.

12.7 Indications for the reagents going bad:

Substrate solution should be abandoned if it turns colors.

The reagents may be turn bad if the absorbance value (450/630nm) of the zero standard is less than 0.5(A450nm＜0.5).

12.8 The coloration reaction need 15min after the addition of solution A and solution B; But you can prolong the incubation time if the color is too light to be determined., never exceed 30min,On the contrary, shorten the incubation time properly.

12.9 The optimal reaction temperature is 25℃. Higher or lower temperature will lead to the changes of sensitivity and absorbance values.

13. Storage condition and storage period

Storage condition: 2-8℃.

Storage period: 6 months.