Competitive Enzyme Immunoassay Kit for

Quantitative Analysis of Tetracyclines

1. Background

It has been dictated in EU No.675/92 Decree(released in March 18th,1992 ) that the total tetracyclines detected in muscle or milk must not exceed 100ppb, and the total tetracyclines in kidney, liver and eggs must not exceed 600ppb, 300ppb and 200ppb.

The determination of tetracycline residue is carried out with HPLC.Sample preparation using this method is time consuming and need expensive experimental equipments. This kit is quite accurate and sensitive, which requires simple operation techniques, less time in detection and can detect more samples.

2. Test Principle

This kit is based on indirect-competitive ELISA technology. The microtiter wells are coated with coupling antigen. Tetracycline residue in the sample competes with the antigen coated on the microtiter plate for the antibody. After the addition of enzyme labeled anti-antibody, TMB substrate is used to show the color. Absorbance of the sample is negatively related to the tetracycline reside in it, after comparing with the Standard Curve, multiplied by the dilution multiple, Tetracycline residue quantity in the sample can be calculated.

3. Applications

This kit can be used in quantitative and qualitative analysis of tetracycline residue in animal tissue (muscle, liver and aquatic product) and honey.

4. Cross-reactions

tetracycline………………………………………………100%

aureomycin…………………………………… about 100%

terramycin……………………………………about 58%

doxycycline……………………………………about 76%

minocycline……………………………………about 76%

demeclocycline……………………………… …about 86%

bristacin…………………………………… …about 78%

Others…………………………………………less than 1%

5. Materials Required

5.1 Equipments:

┅┅Microtiter plate spectrophotometer (450nm/630nm)

┅┅Rotary evaporator or nitrogen drying instruments

┅┅Homogenizer or stamocher

┅┅Shaker

┅┅Gyroscope  or vortex mixer

┅┅Centrifuge

┅┅Analytical balance (inductance: 0.01g)

┅┅Graduated pipette: 10ml

┅┅Rubber pipette bulb

┅┅volumetric flask：100ml、500ml、1L

┅┅glass test rube：10ml

┅┅Polystyrene Centrifuge tubes:2ml、50ml

┅┅Micropipettes: 20(l～200(l、100(l～1000(l

250(l -multiple

5.2 Reagents:

----RIDA C18 column

--- Methanol(AR)

----Citric acid monohydrate

----disodium hydrogen phosphate 12-hydrate(AR)

---- potassium dihydrogen phosphate dihydrate (AR)

---- trichloroacetic acid(AR)

---- sodium hydroxide(AR)

----sodium chloride (AR)

----potassium chloride (AR)

----EDTA-Sodium salt

----Oxalic acid

----deionized water

6. Kit Components

1、Microtiter plate with 96 wells coated with coupling antigen

2、Standard solutions(6 bottles×0.5ml/bottle)

0ppb, 3ppb, 6ppb, 12ppb, 24ppb, 60ppb

3、High concentration standard control:(1ml/bottle)  1ppm

4、concentrated enzyme-labeled secondary antibody solution 1.5ml… red cap

5、Antibody working fluid(7ml) ……………………green cap

6、Solution A 7ml…………………………white cap

7、Solution B 7ml ……………………………… red cap

8、Stop solution  7ml …………………………… yellow cap

9、20×Concentrated wash solution 40ml… transparent cap

10、extraction solution  50ml…………………… blue cap

7. Reagents Preparation:

Solution 1: 0.02 M PBS1 buffer solutions (For honey sample)

Weigh 5.36g disodium hydrogen phosphate 12-hydrate , 0.2g potassium dihydrogen phosphate dihydrate,8.0g sodium chloride and 0.2g potassium chloride,dilute with deionized water to 1L.

Solution 2: Mcllvain-buffer (For Tissue sample treat method (a))

Weigh 12.9g Citric acid monohydrate,10.9g disodium hydrogen phosphate 12-hydrate and 37.2g EDTA-Sodium salt,dilute with deionized water to 1L and adjust the pH to 3.8.

Solution 3: wash solution(For Tissue sample treat method(a) ), methanol solution with 20mM Oxalic acid(=1.8g/L)

Weigh 1.8g Oxalic acid ,dilute with methanol to 1L.

Solution 4: 3% trichloroacetic acid solution (For Tissue sample treat method (b)),

Weigh 3.0g trichloroacetic acid, dilute with deionized water to 100ml.

Solution 5: 1M sodium hydroxide solution (For Tissue sample treat method (b)),

Weigh 4.0g sodium hydroxid,dilute with deionized water to 100ml.

Solution 6: Standard sample solution

Dilute each bottle of concentrated standard sample solution with extraction solution in the volume ratio of 1:9. This standard sample solution can be conserved for 1 day at 4℃.

Solution 7: Wash solution

Dilute the 20×concentrated wash solution with deionized water in the volume ration of 1:19, which will be used to wash the plates. This diluted solution can be conserved for 1 month at 4℃.

8. Sample Preparations

8.1 Notice and precautions for the users before operation:

(a) Please use one-off tips in the process of experiment, and change the tips when absorb different reagent.

(b) Make sure that all experimental instruments are clean.

(c) untreated tissue samples should be stored in freeze.

(d) treated samples should be tested immediately, which can not be conserved at room temperature for more than 4 hours.

(e) Make sure to shake completely when dissolving honey samples with 0.02 M PBS1 buffer solution.

(f) For tissue samples prepared with Method(a) ,after extract twice, merge the extraction slution.Make sure the solution after filter is enough for the column.

(g) 1M NaOH solution mentioned in Method (b)  is used for neutralising the trichloroacetic acid in the sample,remember not to add it after adding the extraction solution.

8.2 Honey

┅┅Weigh 1.0±0.05g honey sample to a 50ml polystyrene centrifuge tube.

┅┅Add 10ml 0.02M PBS1 solution(see solution 1),shake until the honey completely dissolved with shaker.

----homogenize with ultrasonic for 5min.

----Completely mix, and shake for 5min with shaker.

----Take 50 (l prepared solution for sample.

8.3 Tissue (muscle, liver and aquatic product):

Pre-treatment Method(a) :

┅┅Homogenize tissue sample with homogenizer or stamocher.

----Weigh 2.0±0.05g homogenate to a 50ml polystyrene centrifuge tube, add 8ml Mcllvain-buffer (see solution 2), and shake with shaker for 30min.

----Centrifuge for 10min, at least 3000g, at 10℃

----Take 5ml supernatant liquid (never take some suspend suspending impurity).

Filtrate purified with RIDA C18 column follow the next few steps:

----Wash with 3ml methanol;

----Wash with 2ml deionized water;

----Transfer 5ml filtrate into the column;

----Wash the column with 2ml deionized water.

----Carefully extrude all the liquid;

----Wash the sample with 1ml wash solution (see solution 3), keep the flow speed at 15 drops per min, and then collect all the liquid into a glass test tube.

----Dilute the sample solution with extraction solution (see Kit Components 10) in the volume ratio of 1:9.

----Take 50(l of the prepared solution for assay.

Pre-treatment Method(b)

┅┅Homogenize the tissue sample with homogenizer or stamocher;

----Weigh 1.0±0.05g of the homogenate into a 50ml polystyrene centrifuge tube, add 5ml 3% trichloroacetic acid solution (see solution 4), and shake for 5min with shaker to mix the solution completely.

----Centrifuge for 5min, at least 3000g, at 10℃;

----Transfer 400(l of the supernatant liquid to a 2ml polystyrene centrifuge tube, and then add 40(l 1M sodium hydroxide solution (see solution 5), make it completely mixed.

----Add 360(l extraction solution, whorl it for 30s with gyroscope  or vortex mixer to mix it completely.

----Take 50(l of the prepared solution for assay.

9. Assay process

9.1 Notice before assay:

9.1.1 Make sure all reagents and microwells are all at room temperature (20-25℃).

9.1.2 Return all the rest reagents to 2～8℃ immediately after used.

9.1.3 Washing the microwells correctly is an important step in the process of assay; it is the vital factor to the repetitiveness of the ELISA analysis.

9.1.4 Avoid the light and cover the microwells during incubation.

9.2 Assay Steps:

9.2.1 Take all reagents out at room temperature (20-25℃) for more than 30min, homogenize before use.

9.2.2 Get the microwells needed out and return the rest into the zip-lock bag at 2-8℃ immediately.

9.2.3 The diluted wash solution should be rewarmed to be at room temperature before use.

9.2.4 Number: Number every microwell position and all standards and samples should be run in duplicate. Record the standards and samples positions.

9.2.5 Add standard solution/sample: Add 20 µl of standard solution (see solution 6) or prepared sample to corresponding wells. Add 50µl antibody solution. Mix gently by shaking the plate manually and incubate for 60min at 25℃ with cover.

9.2.6 Wash: Remove the cover gently and pure the liquid out of the wells and rinse the microwells with 250µl diluted wash solution (solution 7) at interval of 10s for 4~5 times. Absorb the residual water with absorbent paper (the rest air bubble can be eliminated with unused tip).

9.2.7 Dilution of enzyme-labeled secondary antibody solution: dilute the enzyme-labeled secondary antibody solution with wash solution in the volume ratio of 1:9 according to the quantity you need. (Notice: the diluted solution can not be stored, you have to dilute again when use)

9.2.8 Add enzyme-labeled secondary antibody solution: Add enzyme-labeled secondary antibody 100(l to each well, Mix gently by shaking the plate manually and incubate for 30min at 25℃ with cover. Repeat the wash step again.

9.2.8 Coloration: Add 50µl solution A and 50µl solution B to each well. Mix gently by shaking the plate manually and incubate for 15~20min at 25℃ with cover(see 12.8).

9.2.9 Measure: Add 50µl the stop solution to each well. Mix gently by shaking the plate manually and measure the absorbance at 450nm against an air blank (It’s suggested measure with the dual-wavelength of 450/630nm. Read the result within 5min after addition of stop solution. ) (We can also measure by sight without stop solution in short of the ELIASA instrument)

10. Results

There are 2 different methods to determinate the results. Method 1 leads to a round estimation, and method 2 leads to definite quantity estimation. (Please notice: the absorption is inversely proportional to the Tetracycline concentration in the sample.)

10.1 Round estimation

We can get the range of different strengths from the compare of average absorption and standards by sight. For example, the absorption of sample 1 is 0.321, sample 2 is 0.982, and the absorptions of Tetracycline standard solutions: 1.631 (0ppb); 1.330(0.3ppb); 1.130(0.6ppb); 0.854(1.2ppb); 0.443(2.4ppb); 0.242(6.0ppb). So we can say the strength of diluted sample 1 is between 2.4ppb and 6.0ppb; and diluted sample 2 between 0.6 ppb to 1.2 ppb. In order to obtain the Tetracycline actually contained in a sample, the diluted sample results must be further multiplied by the corresponding dilution multiple.

10.2 Definite quantity estimation

(1) The mean values of the absorbance values obtained for the standards and the samples are divided by the absorbance value of the first standard (zero standard ) and multiplied by 100%. The zero standard is thus made equal to 100% and the absorbance values are quoted in percentages.

B

Absorbance (%) = —— ×100%

B0

B ——absorbance standard (or sample)

B0 ——absorbance zero standard

(2) To draw a standard curve: Take the absorbance value of standards as y-axis, semi logarithmic of the concentration of the Tetracycline standards solution (ppb) as x-axis.

--- The Tetracycline concentration of each sample (ppb), which can be read from the calibration curve, is multiplied by the corresponding dilution multiple of each sample followed, and the actual concentration of sample is obtained.

For evaluation of the ELISA kits special software has been developed for exact and rapid analysis. The analysis software can be ordered on request.

Dilution multiple of samples：

Honey：10

Tissue samples treated with Method(a) ：10

Tissue samples treated with Method(b) ：10

11. Sensitivity, accuracy and precision

Test Sensitivity：0.3ppb

Detection limit：

Detection limit of Tissue sample

Tetracycline……………………………………3ppb

aureomycin………………………………………about 3ppb

terramycin………………………………………about 5ppb

doxycycline……………………………………about 4ppb

minocycline……………………………………about 4ppb

demeclocycline………………………………about 4ppb

bristacin………………………………about 4ppb

Detection limit of Honey sample

Tetracycline…………………………3ppb

aureomycin………………………about 3ppb

terramycin…………………………about 5ppb

doxycycline…………………………about 4ppb

minocycline…………………………about 4ppb

demeclocycline……………………about 4ppb

bristacin………………………….…about 4ppb

Accuracy:

Tissue ………………………………………………80±15%

Honey ………………………………………………95±15%

Precision:

Variation coefficient of the ELISA kit is less than 10%.

12. Notice

12.1 The mean values of the absorbance values obtained for the standards and the samples will be reduced if the reagents and samples have not been regulated to room temperature (20-25℃).

12.2 Do not allow microwells to dry between steps to avoid unsuccessful repetitiveness and operate the next step immediately after tap the microwells holder.

12.3. Homogenize each reagent before using.

12.4. Keep your skin away from the stop solution for it is the 2M H2SO4 solution.

12.5 Don’t use the kits out of date. Don’t exchange the reagents of different batches, for it will drop the sensitivity.

12.6 Storage condition:

Keep the ELISA kits at 2-8℃,do not freeze. Seal rest microwell plates Avoid straight sunlight during all incubations. Covering the microtiter plates is recommended.

12.7 Indications for the reagents going bad:

Substrate solution should be abandoned if it turns colors.

The reagents may be turn bad if the absorbance value (450/630nm) of the zero standard is less than 0.5(A450nm＜0.5).

12.8 The coloration reaction need 15-20min after the addition of solution A and solution B; But you can prolong the incubation time ranges from 20min to more if the color is too light to be determined., never exceed 30min,On the contrary, shorten the incubation time properly.

12.9 The optimal reaction temperature is 25℃. Higher or lower temperature will lead to the changes of sensitivity and absorbance values.

13. Storage condition and storage period

Storage condition: 2-8℃.

Storage period: 6 months.