Competitive Enzyme Immunoassay Kit for

Quantitative Analysis of Norethindrone

1. Background

Norethindrone is an anabolic steroid hormone, which is broadly applied in promoting animal growth. It is now being banned or prohibited in developed countries such as US and the EU for obvious carcinogenicity.  Document No. 235 of Ministry of Agriculture prescribes that the residual of norethindrone in animal-derived food must not be detected. GC is the common approach to detect this drug, but it is restricted for the complicated operation and high expense. While this kit is a new product for drug residual detection based on ELISA technology, which only costs 1.5 hours in each operation and can considerably minimize operation errors and work intensity

2. Test Principle

This kit is based on indirect-competitive ELISA technology. The microtiter wells are coated with coupling antigen. Norethindrone residual in the sample competes with the antigen coated on the microtiter plate for the antibody. After the addition of horseradish peroxidase labeled anti-antibody, tetramethyl-benzidine (TMB) substrate is used to show the color. Absorbance of the sample is negatively related to the Norethindrone residual in it, after comparing with the Standard Curve, multiplied by the dilution multiples, Norethindrone residual quantity in the sample can be calculated.

3. Applications

This kit can be used in quantitative and qualitative analysis of Norethindrone residual in animal tissue (muscle, liver, fish and shrimp) feed and urine, etc.

4. Cross-reactions

Norethindrone…………………………………… 100%

megaesirol ……………………………………… 0.1%

medroxyprogesterone acetate………………… 0.2%

trenbolone…………………………………………0.3%

medroxyprogesterone……………………………0.25%

progesterone……………………………….……﹤0.1%

diethylstilbestrol…………………………………﹤0.1%

estradiol……………………………………………0.3%

levonorgestrel…………………………………………2%

Norgestrel…………………………………………100%

5. Materials Required

5.1 Equipments:

┅┅Microtiter plate spectrophotometer (450nm/630nm)

┅┅Rotary evaporator or nitrogen drying instruments

┅┅Homogenizer or stamocher

┅┅Shaker

┅┅Gyroscope  or vortex mixer

┅┅Centrifuge

┅┅Analytical balance (inductance: 0.01g)

┅┅Graduated pipette: 10ml

┅┅Rubber pipette bulb

┅┅ Volumetric flask : 100ml , 500ml ;

┅┅ Glass test tube : 15 ml ;

┅┅Polystyrene centrifuge tube: 50ml

┅┅Micropipettes: 20(l～200(l、100(l～1000(l

250(l-multipipette

Reagents：

----sodium nitroferricyanide dihydrate (AR) (for tissue samples)

---- Zinc sulphate heptahydrate  (AR) (for tissue samples)

---- Acetonitrile (AR) (for tissue samples)

----Sodium hydroxide (AR)

----deionized water ;

6. Kit Components

1、Microtiter plate with 96 wells coated with coupling antigen

2、standard solutions(6 bottles:1ml/bottle)

0ppb,  0.1ppb,  0.3ppb,  0.9ppb,  2.7ppb,  8.1ppb

3、High concentration standard control: (1ml/bottle)  100ppb

4Enzyme-labeled secondary antibody solution (12ml)…red cap

5、Antibody solution ( 7ml) ……………………green cap

6、Solution A (7ml)……………………………white cap

7、Solution B (7ml)…………………………red cap

8、Stop solution (7ml) …………………………yellow cap

9、20×concentrated Wash solution(40ml) ………transparent cap

10、2×Concentrated extraction solution (50ml)………blue cap

7. Reagents Preparation:

solution 1: solution C (for tissue samples)

0.36M sodium nitroferricyanide dihydrate solution (Na2Fe(CN)5·NO·2H2O)

Weigh 10.7g sodium nitroferricyanide dihydrate, dissolve with deionized water and dilute to 100ml;

solution 2: solution D (for tissue samples)

1M Zinc sulphate solution (ZnSO4·7H2O)

Weigh 28.8g zinc sulphate heptahydrate ,dissolve with deionized water and dilute to 100ml;

solution 3: 0.1M sodium hydroxide solution

Dissolve 0.4g sodium hydroxide with deionized water and dilute to 100ml;

solution 4: acetonitrile-0.1M sodium hydroxide solution (for tissue samples)

Weigh 84ml acetonitrile mix with 16ml 0.1M sodium hydroxide solution ;

solution 5: acetonitrile -0.1M sodium hydroxide solution (for feed samples)

Weigh  80ml acetonitrile mix with 20ml 0.1M sodium hydroxide solution ;

solution 6: Extraction solution

Dilute the concentrated extraction solution with deionized water in the volume ratio of 1: 1, which will be used for sample extraction. This solution can be conserved for a month at 4℃.

solution 7: Wash solution

Dilute the concentrated wash solution with deionized water in the volume ratio of 1:19(or according to the requirement), which will be used to rinse the plates. The diluted wash solution can be conserved for a month at 4℃.

8. Sample Preparations

8.1 Notice and precautions for the users before operation:

(a) Please use one-off tips in the process of experiment, and change the tips when absorbing different reagent.

(b) Make sure that all experimental instruments are clean.

(c) extracted solution of tissue and feed samples can be conserved at room temperature for 12 hours.

8.2  Tissue samples (muscle and liver of chicken and pig, shrimp and fish ,etc.)

---- Homogenize the samples with homogenizer or stamocher;

----Weigh 2.0±0.05g of the homogenate into a 50ml polystyrene centrifuge tube, then add 7ml acetonitrile -0.1M sodium hydroxide solution (for tissue samples)(see solution 4), add  500µl  of Solution C and solution D (see solution 1 and solution 2) , shake completely for 10min with shaker, centrifuge at room temperature (20-25℃) for 10min, at least 3000g;

----Take 1ml of the supernate into a 10ml clean glass tube, dry with 50～60℃ nitrogen gas flow ;

----Dissolve the dry leftover with 1ml extraction solution (see solution 6) ;

----Dilute the extracted sample solution with extraction solution in the volume ration of 1:4 (100µl extracted sample solution + 400µl extraction solution ).

---- Take 50(l of the prepared solution for assay ;

8.3 Feed sample:

----Weigh 1.0±0.05g feed sample into a 50ml polystyrene  centrifuge tube, add 10ml acetonitrile-0.1M sodium hydroxide solution (for feed sample) (see solution 5), shake for 5min with shaker, then centrifuge at room temperature(20-25℃) for 5min, at least 3000g ;

----Take 0.5ml of the supernate into a 10ml clean glass test tube, add 0.5ml 0.1M sodium hydroxide solution (see solution 3), whorl for 20s, then add 5ml chloroform, shake for 5min with shaker, then centrifuge at room temperature(20-25℃) for 5min, at least 3000g ;

----Take 1ml of the substrate transparent liquid into a 10ml clean glass test tube, dry with 50～60℃ nitrogen gas flow;

---- Dissolve the dry leftover with 1ml extraction solution (see solution 6) ;

----Dilute the extracted sample solution with extraction solution in the volume ration of 1:4 (100µl extracted sample solution + 400µl extraction solution ).

---- Take 50(l of the prepared solution for assay ;

9. Assay process

9.1 Notice before assay:

9.1.1 Make sure all reagents and microwells are all at room temperature (20-25℃).

9.1.2 Return all the rest reagents to 2～8℃ immediately after used.

9.1.3 Washing the microwells correctly is an important step in the process of assay; it is the vital factor to the reproducibility of the ELISA analysis.

9.1.4 Avoid the light and cover the microwells during incubation.

9.2 Assay Steps:

9.2.1 Take all reagents out at room temperature (20-25℃) for more than 30min, homogenize before use.

9.2.2 Get the microwells needed out and return the rest into the zip-lock bag at 2-8℃ immediately.

9.2.3 The diluted wash solution should be rewarmed to be at room temperature before use.

9.2.4 Number: Numbered every microwell positions and all standards and samples should be run in duplicate. Record the standards and samples positions.

9.2.5 Add standard solution/sample: Add 50 µl of standard solution or prepared sample to corresponding wells. Add 50µl antibody solution. Mix gently by rocking the plate manually and incubate for 30min at 37℃ with cover.

9.2.6 Wash: Remove the cover gently and pure the liquid out of the wells and rinse the microwells with 250µl diluted wash solution (solution 7) at interval of 10s for 4~5 times. Absorb the residual water with absorbent paper (the rest air bubble can be eliminated with unused tip).

9.2.7 Enzyme-labeled secondary antibody: Add Enzyme-labeled secondary antibody solution 100(l to each well, Mix gently by rocking the plate manually and incubate for 30min at 37℃ with cover. Then repeat the wash step.

9.2.8 Coloration: Add 50µl solution A and 50µl solution B to each well. Mix gently by rocking the plate manually and incubate for 15min at 37℃ with cover(see 12.8).

9.2.9 Measure: Add 50µl the stop solution to each well. Mix gently by rocking the plate manually and measure the absorbance at 450nm against an air blank (It’s suggested measure with the dual-wavelength of 450/630nm. Read the result within 5min after addition of stop solution. ) (We can also measure by sight without stop solution in short of the ELIASA instrument)

10. Results

There are 2 different methods to determinate the results. Method 1 leads to a round estimation, and method 2 leads to definite quantity estimation. (Please notice: the absorption is inversely proportional to the Norethindrone concentration in the sample.)

10.1 Round estimation

We can get the range of different strengths from the compare of average absorption and standards by sight. For example, the absorption of sample 1 is 0.310, sample 2 is 0.820, and the absorptions of Norethindrone standard solutions: 2.020 (0ppb); 1.512(0.1ppb); 0.957(0.3ppb); 0.506(0.9ppb); 0.172(2.7ppb); 0.08(8.1ppb). So we can say the strength of diluted sample 1 is between 0.9ppb and 2.7ppb; and diluted sample 2 between 0.3 ppb to 0.9ppb. In order to obtain the Norethindrone actually contained in a sample, the diluted sample results must be further multiplied by the corresponding dilution factor.

10.2 Definite quantity estimation

(1) The mean values of the absorbance values obtained for the standards and the samples are divided by the absorbance value of the first standard (zero standard ) and multiplied by 100%. The zero standard is thus made equal to 100% and the absorbance values are quoted in percentages.

B

Absorbance (%) = —— ×100%

B0

B ——absorbance standard (or sample)

B0 ——absorbance zero standard

(2) To draw a standard curve: Take the absorbance value of standards as y-axis, semi logarithmic of the concentration of the Norethindrone standards solution (ppb) as x-axis.

--- The Norethindrone concentration of each sample (ppb), which can be read from the calibration curve, is multiplied by the corresponding Dilution factor of each sample followed, and the actual concentration of sample is obtained.

Please notice:

For evaluation of the ELISA kits special software has been developed for exact and rapid analysis. The analysis software can be ordered on request.

Sample dilution multiple:

Tissue sample：20

Feed sample：500

11. Sensitivity, accuracy and precision

Test Sensitivity:0.1ppb

Detection limit:

Tissue…………………………………………1ppb

Feed………………………………………50ppb

Accuracy:

Tissue…………………………………………… 80±15%

Feed…………………………………………… 90±20%

Precision:

Variation coefficient of the ELISA kit is less than 10%.

12. Notice

12.1 The mean values of the absorbance values obtained for the standards and the samples will be reduced if the reagents and samples have not been regulated to room temperature (20-25℃).

12.2 Do not allow microwells to dry between steps to avoid unsuccessful reproducibility and operate the next step immediately after tap the microwells holder.

12.3. Homogenize each reagent before using.

12.4. Keep your skin away from the stop solution for it is the 2M H2SO4 solution.

12.5 Don’t use the kits out of date. Don’t exchange the reagents of different batches, or else it will drop the sensitivity.

12.6 Storage condition:

Keep the ELISA kits at 2-8℃, do not freeze. Seal rest microwell plates, Avoid straight sunlight during all incubations. Covering the microtiter plates is recommended.

12.7 Indications for the reagents going bad:

Substrate solution should be abandoned if it turns colors.

The reagents may be turn bad if the absorbance value (450/630nm) of the zero standard is less than 0.5(A450nm＜0.5).

12.8 The coloration reaction need 15min after the addition of solution A and solution B; But you can prolong the incubation time ranges from 20min to more if the color is too light to be determined., never exceed 30min, On the contrary, shorten the incubation time properly.

12.9 The optimal reaction temperature is 37℃. Higher or lower temperature will lead to the changes of sensitivity and absorbance values.

13. Storage condition and storage period

Storage condition: 2-8℃.

Storage period: 6 months.