EnzyChromTM L-Lactate Assay Kit (ECLC-100)

Ultrasensitive Colorimetric Determination of L-Lactate at 565 nm

DESCRIPTION

Lactate is generated by lactate dehydrogenase (LDH) under hypoxic or anaerobic conditions. Monitoring lactate levels is, therefore, a good indicator of the balance between tissue oxygen demand and utilization and is useful when studying cellular and animal physiology.

Simple, direct and automation-ready procedures for measuring lactate concentration are very desirable. BioAssay Systems' EnzyChromTM lactate assay kit is based on lactate dehydrogenase catalyzed oxidation of lactate, in which the formed NADH is coupled to the formazan (MTT)/phenazine methosulfate (PMS) Reagent. The intensity of the product color, measured at 565 nm, is proportionate to the lactate concentration in the sample.

APPLICATIONS

Direct Assays: lactate in serum, plasma, and cell media samples.

KEY FEATURES

Sensitive and accurate. Detection limit of 0.05 mM and linearity up to 2 mM L-Lactate in 96-well plate assay. For cell culture samples containing phenol red: detection limit of 0.1 mM and linearity up to 1 mM L-Lactate in 96-well plate assay.

Convenient. The procedure involves adding a single working reagent, and reading the optical density at time zero and at 20 min. Room temperature assay. No 37°C heater is needed.

High-throughput. Can be readily automated as a high-throughput 96- well plate assay for thousands of samples per day.

KIT CONTENTS (100 tests in 96-well plates)

Assay Buffer: 6 mL	NAD Solution: 1 mL

PMS Solution: 1.5 mL	MTT Solution: 1.5 mL

Enzyme: 120 RL L-LDH Standard: 1.0 mL 20 mM L-Lactate

Storage conditions. Store all reagents at -20°C. Shelf life of at least 6 months (see expiry dates on labels).

Precautions: reagents are for research use only. Normal precautions for laboratory reagents should be exercised while using the reagents. Please refer to Material Safety Data Sheet for detailed information.

PROCEDURES

Important: this assay is based on an enzyme-catalyzed kinetic reaction. Addition of Working Reagent should be quick and mixing should be brief but thorough. Use of a multi-channel pipettor is recommended.

1. Standard Curve. Prepare 1000 RL 2.0 mM L-lactate Premix by mixing 100 RL 20 mM Standard and 900 RL distilled water. For cell culture samples containing phenol red, prepare 1000 RL 1.0 mM lactate Premix by mixing 50 RL 20 mM Standard and 950 RL culture media. Dilute standard as follows. Transfer 20 RL standards into wells of a clear bottom 96-well plate.

| No | Premix + H2O or Media | Vol (RL) | L-Lactate (mM) |
| --- | --- | --- | --- |
| 1 | 100RL + 0RL | 100 | 2.0 or 1.0 |
| 2 | 80RL + 20RL | 100 | 1.6 or 0.8 |
| 3 | 60RL + 40RL | 100 | 1.2 or 0.6 |
| 4 | 40RL + 60RL | 100 | 0.8 or 0.4 |
| 5 | 30RL + 70RL | 100 | 0.6 or 0.3 |
| 6 | 20RL + 80RL | 100 | 0.4 or 0.2 |
| 7 | 10RL + 90RL | 100 | 0.2 or 0.1 |
| 8 | 0RL + 100RL | 100 | 0 |

Samples. Add 20 RL sample per well in separate wells. For samples with potential endogenous enzyme activity (i.e. serum, plasma, tissue extracts), two reactions should be run: one with added Enzyme and a No Enzyme control. Serum and Plasma should be diluted at least 2x with dH2O.

2.	Reagent Preparation. Spin the Enzyme tube briefly before pipetting. For each reaction well, prepare Working Reagent by mixing 50 RL Assay Buffer, 1 RL Enzyme, 10 RL NAD, 14 RL PMS and 14 RL MTT. Fresh reconstitution is recommended. For the No Enzyme control, the Working Reagent includes 50 RL Assay Buffer, 10 RL NAD, 14 RL PMS and 14 RL MTT.

3.	Reaction. Add 80 RL Working Reagent per reaction well quickly. Tap plate to mix briefly and thoroughly.

4.	Read optical density (OD0) for time “zero” at 565 nm (520-600nm) and OD20 after a 20-min incubation at room temperature.

5.	Calculation. Subtract OD0 from OD20 for the standard and sample wells. Use the DOD values to determine the sample L-lactate concentration from the standard curve. For samples requiring a No Enzyme control, subtract the DODNoEnz value from the DODSample and use this DDOD value to determine the sample L-lactate concentration from the standard curve.

Note: if the sample OD value is higher than OD for 2 mM L-lactate standard, dilute sample in water and repeat the assay. Multiply the results by the dilution factor.

MATERIALS REQUIRED, BUT NOT PROVIDED

Pipeting (multi-channel) devices. Clear-bottom 96-well plates (e.g. Corning Costar) and plate reader.

GENERAL CONSIDERATIONS

The following substances interfere and should be avoided in sample preparation: ascorbic acid, SDS (>0.2%), sodium azide, NP-40 (>1%) and Tween-20 (>1%).

0.6

0.6

L-Lactate

L-Lactate

0.4

0.4

0.2

0.2

0.0

0.0

0.0	0.5	1.0	1.5	2.0

0.0	0.5	1.0	1.5	2.0

[L-Lactate](mM)

Standard Curve in 96-well plate assay in water.

LITERATURE

[1].	Babson, AL and Babson, SR. (1973) Kinetic Colorimetric Measurement of Serum Lactate Dehydrogenase Activity. Clin Chem. 19(7):766-9.

[2].	Karlsen RL, Norgaard L, Guldbrandsen EB (1981). A rapid method for the determination of urea stable lactate dehydrogenase on the 'Cobas Bio' centrifugal analyser.Scand J Clin Lab Invest. 41(5):513-6.

[3].	Coley HM, Lewandowicz G, Sargent JM, Verrill MW (1997). Chemosensitivity testing of fresh and continuous tumor cell cultures	using	lactate dehydrogenase.Anticancer	Res.
17(1A):231-6.