Influenza (Av.)

H5-HA(Ab)

H5-HA(Ab)

ELISA KIT FOR DETECTION OF ANTIBODIES TO
HEMAGGLUTININ (HA) OF INFLUENZA A VIRUS, H5 STRAIN

ELISA KIT FOR DETECTION OF ANTIBODIES TO
HEMAGGLUTININ (HA) OF INFLUENZA A VIRUS, H5 STRAIN

One-Step Incubation, Competitive Principle

One-Step Incubation, Competitive Principle

INSTRUCTIONS FOR USE, 2005

INSTRUCTIONS FOR USE, 2005

INTENDED USE

This kit is an enzyme-linked immunosorbent assay (ELISA) for qualitative detection of antibodies (anti-HA) to hemagglutinin (HA) of Influenza A viruses, H5 strain (also known as highly pathogenic avian influenza) in human serum or plasma samples. It is intended for clinical identification of infection with the H5 strain of type-A influenza viruses.

SUMMARY

Influenza infection is an acute fever-like virus infection of the respiratory tract. The influenza virus and its toxin can lead to a serious inflammation of the bronchial mucosa and a damage of the tract.

The influenza viruses belong to the family of Orthomyxoviridae that have linear segmented (8) negative ssRNA genome with lipid envelope. Total genome length is 12000-15000 nucleotides (nt). The largest segment is 2300-2500 nt; the second largest is 2300-2500 nt; the third is 2200-2300 nt; the fourth is 1700-1800 nt; the fifth is 1500-1600 nt; the sixth is 1400-1500 nt; the seventh is 1000-1100 nt; the eighth is 800-900 nt. Genome sequence has terminal repeated sequences; repeated at both ends.

The virion envelope could be spherical, or filamentous with diameter of 50-120 nm, or 20 nm and 200-300(-3000) nm long. About 500 spikes are dispersed evenly over all the surface (i.e. hemagglutininesterase (HEF)), or dispersed equally over all the surface, but the various types are in clusters (i.e. hemagglutinin (HA) the major glycoprotein is interposed irregularly by clusters of neuraminidase (NA), with ratio of HA to NA about 4-5 to 1).

The Orthomyxoviridae family is divided into three types : A,B,C.

Type A influenza viruses are further divided into subtypes according to their Hemagglutinin (HA) and Neuraminidase (NA) proteins. Currently 15 (HA) subtypes and 9(NA) subtypes have been identified.

Type B influenza viruses produce less serious disease than does influenza type A and are not categorized as by H or N type as Influenza A is.

Type C influenza viruses were first isolated in 1949 and are not known to be responsible for epidemics

The infection in human mostly results from a droplet infection
and appears as an epidemic which sometimes can be of
pandemic proportions. After an incubation time of 1 - 3 days

the symptoms appear suddenly- fast raise of temperature, often accompanied by shivering, the leading symptom of catarrhal inflammation appears, contributing to the clinical course of painful dry cough, tracheitis, laryngitis and frequent rhinitis and conjunctivitis.

The appearance new Influenza epidemics and pandemics are facilitated by an antigen variability mainly in the HA and NA. In the past century, three major influenza epidemics have occurred:

1918-1919: (Spanish Flu, A (H1N1), 20-50 million deaths worldwide, nearly half were young, healthy adults.

1957-1958: Asian Flu (A (H2N2),1st identified in China Feb. 1957, 70,000 deaths in the United States.

1968-1969: Hong Kong Flu (A (H3N2), 1st detected in Hong Kong early 1968, virus still circulating today.

The 1997 outbreak of H5N1 avian influenza (or bird flu) in humans in Hong Kong caused alarm because it involved highly pathogenic strains of an influenza subtype A to which humans lack immunity. The H5N1 infected 18 humans, 6 of whom died (death rate of about 70 percent). Most of these cases occurred from contact with infected poultry or contaminated surfaces; however, it is thought that a few cases of human-to-human spread of H5N1 have occurred.

Avian influenza is an infectious disease of birds caused by type A strains of the influenza virus. The disease, first identified in Italy more than 100 years ago, now occurs worldwide. Infection triggers a wide spectrum of symptoms in birds, ranging from mild illness to a highly contagious and rapidly fatal disease resulting in severe epidemics. In the H5N1 bird flu in Hong Kong in 1997, patients had developed symptoms of fever, sore throat, cough and, in several of the fatal cases, severe respiratory distress secondary to viral pneumonia. Previously healthy adults and children, and some with chronic medical conditions, were affected.

More recently, outbreaks of avian influenza H5N1 occurred among poultry in eight countries in Asia (Cambodia, China, Indonesia, Japan, Laos, South Korea, Thailand, and Vietnam) during late 2003 and early 2004. At that time, more than 100 million birds in the affected countries either died from the disease or were killed in order to try to control the outbreak. By March 2004, the outbreak was reported to be under control. Beginning in late June 2004, however, new deadly outbreaks of influenza H5N1 among poultry were reported by several countries in Asia (Cambodia, China, Indonesia, Malaysia [first-time reports], Thailand, and Vietnam ). It is believed that these outbreaks are ongoing.

Human infections of avian influenza H5N1 however, have been reported in Cambodia (1case/1death) Thailand (17cases/1 death) and Vietnam (51 cases/ 33deaths) during both of these outbreak periods.

Hemagglutinin (HA) is a surface glycoprotein on Influenza A responsible for binding to N-AcetylNeuraminic Acid (NeuNAc) or commonly Sialic acid on host cell surface receptors. The Influenza viruses form the A virus group have principally similar morphological, chemical and biological features. The differentiation of the types is possible by the different antigenicity of their nucleo- and matrix proteins that have type-specific antigenicity. However, the essential immunodominant antigens and primary targets in diagnosis are the hemagglutinin (HA) and the neuraminidase (NA) antigens. Screening for type-specific anti-HA or anti-NA antibodies has also been proved to be useful method in clinical identification of different influenza strains.

PRINCIPLE OF THE ASSAY

H5-HA (Ab) ELISA kit is based on solid phase, one step incubation competitive principle ELISA. H5 Influenza specific, anti hemagglutinin (anti-HA) antibodies, if present in the sample, compete with monoclonal Influenza H5 antibodies conjugated to horseradish peroxidase (anti-HA”mAb”HRP) for a fixed amount immunological Ag-Ab complex that has been pre-coated inside wells. This complex contains viral lysate of H5 HA, and monoclonal H5 specific anti-HA. When no H5 specific anti-HA presents in the sample, anti-HA”mAb”HRP will be bound with the H5 HA inside the wells and any unbound HRP-Conjugate is removed during washing. Chromogen A and B solutions are added into the wells and during incubation, the colorless Chromogens are hydrolyzed by the bound HRP-Conjugate to a blue colored product. The blue color turns yellow after stopping the reaction with sulfuric acid. No or low color developing suggests the presence of specific antibodies to hemagglutinin of influenza Type A, H5 strain.

Assay principle scheme: Competitive ELISA

Ab(p)-Ag(p)+Ab(s)+(Ab)ENZ -. [Ab(p)-Ag(p)–Ab(s)	] -. No color	( + )

Ab(p)-Ag(p)	+(Ab)ENZ -. [Ab(p)-Ag(p) – (Ab)ENZ ] -' Blue -.Yellow Color ( - )

Incubation	Immobilized Complex Coloring	Results

30 min.	30 min.

Ab(p)-Ag(p) –pre-coated H5 anti-HA”mAb” –HA complex; Ab(s)–H5 HA specific antibodies in sample;

(Ab)ENZ–HRP conjugated H5 anti-HA”mAb”;

COMPONENTS

·

96 Tests

96 Tests

MICROWELL PLATE	1plate

Blank microwell strips fixed on white strip holder. The plate is sealed in aluminum pouch with desiccant.

8X12 or 12X8-well strips per plate.

Each well contains immunological complex of monoclonal antibodies (anti-HA”mAb”)and H5 HA.

The microwell strips can be broken to be used separately. Place unused wells in the plastic sealable storage bag together with the desiccant and return to 2~8°C.

·	NEGATIVE CONTROL	1vial

Yellowish liquid filled in a vial with green screw cap

0.5ml per vial.

Protein-stabilized buffer tested non-reactive for antibodies to HA of H5 influenza-A viruses.

Preservatives: 0.1% ProClin 300.

Ready to use as supplied.

Once open, stable for one month at 2-8°C.

·	POSITIVE CONTROL	1vial

Red-colored color liquid filled in a vial with red screw cap. 0.5ml per vial.

Protein-stabilized buffer tested reactive for antibodies to HA of H5 influenza-A viruses. Preservatives: 0.1% ProClin 300. Ready to use as supplied.

Once open, stable for one month at 2-8°C.

·	HRP-CONJUGATE REAGENT	1vial

Red-colored liquid filled in a white dropper vial with red screw cap.

6ml per vial. Horseradish peroxidase-conjugated H5 specific anti-HA”mAb” (monoclonal).

Ready to use as supplied.

Once open, stable for one month at 2-8°C.

·	STOCK WASH BUFFER	1 bottle

Colorless liquid filled in a clear bottle with white screw cap. 50ml per bottle.

PH 7.4 20 × PBS (Containing Tween-20 as a detergent). DILUTE BEFORE USE -The concentration must be diluted 1 to 20 with distilled/deionized water before use. Once diluted, stable for one week at room temperature

or for two weeks at 2-8°C.

·	CHROMOGEN SOLUTION A	1vial

Colorless liquid filled in a white dropper vial with green screw cap. 6ml per vial.

Urea peroxide solution.

Ready to use as supplied.

Once open, stable for one month at 2-8°C.

·	CHROMOGEN SOLUTION B	1vial

Colorless liquid filled in a black dropper vial with black screw cap. 6ml per vial.

TMB solution. Tetramethylbenzidine dissolved in citric acid. Ready to use as supplied.

Once open, stable for one month at 2-8°C.

·	STOP SOLUTION	1vial

Colorless liquid filled in a white dropper vial with yellow screw cap 6ml per vial.

Diluted sulfuric acid solution (2.0M H2SO4).

·	PLASTIC SEALABLE BAG	1unit
For enclosing the strips not in use.

·	CARDBOARD PLATE COVER	1sheet

To cover the plates during incubation and prevent evaporation or contamination of the wells.

·	PACKAGE INSERTS	1copy

ADDITIONAL MATERIALS AND INSTRUMENTS REQUIRED
BUT NOT PROVIDED

1.	Freshly distilled or deionized water.

2.	Disposable gloves and timer.

3.	Appropriate waste containers for potentially contaminated materials.

4.	Disposable V-shaped troughs.

5.	Dispensing system and/or pipette (single or multichannel), disposable pipette tips

6.	Absorbent tissue or clean towel.

7.	Dry incubator or water bath, 37± 0.5°C .

8.	Microshaker for dissolving and mixing conjugate with samples.

9.	Microwell plate reader, single wavelength 450nm or dual wavelength 450nm and 630nm.

10.	Microwell aspiration/wash system.

SPECIMEN COLLECTION, TRANSPORTATION AND
STORAGE

1. Sample Collection: Either fresh serum or plasma samples can be used for this assay. Blood collected by venipuncture should be allowed to clot naturally and completely – the serum/plasma must be separated from the clot as early as possible as to avoid hemolysis of the RBC. Care should be taken to ensure that the serum samples are clear and not contaminated by microorganisms. Any visible particulate matters in the

sample should be removed by centrifugation at 3000 RPM for at least 20 minutes at room temperature, or by filtration on 0.22u filters. Plasma samples collected into EDTA, sodium citrate or heparin may be tested, but highly lipaemic, icteric, or hemolized samples should not be used as they could give erroneous results in the assay. Do not heat inactivate samples. This can cause sample deterioration.

2. Transportation and Storage: Store samples at 2-8°C. Samples not required for assaying within 3 days should be stored frozen (-20°C or lower).Multiple freeze-thaw cycles should be avoided. For shipment, samples should be packaged and labeled in accordance with the existing local and international regulations for transport of clinical samples and ethological agents.

SPECIAL INSTRUCTIONS FOR WASHING

1.	A good washing procedure is essential to obtain correct and precise analytical data.

2.	It is therefore recommended to use a good quality ELISA microplate washer, maintained at the best level of washing performances. In general, no less than 5 automatic washing cycles with dispensing of 350-400pl/well, are sufficient to avoid false positive reactions and high background (all wells turn yellow).

3.	To avoid cross-contaminations of the plate with sample or HRP-conjugate, after incubation do not discard the content of the wells, but allow the plate washer to aspirate it automatically.

4.	Anyway, we recommend calibrating the washing system on the kit itself in order to match the declared analytical performances. Assure that the microplate washer’s liquid dispensing channels are not blocked or contaminated, and sufficient volume of Wash buffer is dispensed each time into the wells.

5.	In case of manual washing, we suggest to perform at least 5cycles, dispensing 350-400pl/well and aspirating the liquid for 5times. If poor results (high background) are observed, increase the washing cycles or soaking time per well.

1.	In any case, the liquid aspirated out the strips should be treated with a sodium hypochlorite solution(final concentration of 2.5%) for 24 hours, before liquids are disposed in an appropriate way.

2.	The concentrated Washing solution should be diluted 1 to 20 before use. For one plate, mix 50 ml of the concentrate with 950ml of water for a final volume of 1000ml diluted Wash Buffer. If less than a whole plate is used, prepare the proportional volume of solution.

STORAGE AND STABILITY

The components of the kit will remain stable through the expiration date indicated on the label and package when stored between 2-8 °C , do not freeze. To assure maximum performance of H5-HA(Ab) ELISA kit, during storage, protect the reagents from contamination with microorganism or chemicals.

PRECAUTIONS AND SAFETY

This kit is intended FOR IN VITRO USE ONLY

FOR PROFESSIONAL USE ONLY

The ELISA assay is a time and temperature sensitive method. To avoid incorrect result, strictly follow the test procedure steps and do not modify them.

1.	Do not exchange reagents from different lots, or use reagents from other commercially available kits. The components of the kit are precisely matched as to achieve optimal performance during testing.

2.	Make sure that all reagents are within the validity indicated on the kit box and are of the same lot. Never use reagents beyond the expiry date stated on reagents labels or on the kit box.

3.	CAUTION - CRITICAL STEP: Allow the reagents and samples to stabilize at room temperature (18-30°C) before use. Shake reagent gently before, and return to 2-8°C immediately after use.

4.	Use only sufficient volume of sample as indicated in the procedure steps. Failure to do so, may cause in low sensitivity of the assay.

5.	Do not touch the bottom exterior of the wells; fingerprints or scratches may interfere with microwell reading.

6.	When reading the results, ensure that the plate bottom is dry and there are no air-bubbles inside the wells.

7.	Never allow the microplate wells to dry after the washing step. Immediately proceed to the next step. Avoid the formation of air-bubbles when adding the reagents.

8.	Avoid assay steps long time interruptions. Assure same working conditions for all wells.

9.	Calibrate the pipette frequently to assure the accuracy of samples/reagents dispensing. Always use different disposal pipette tips for each specimen and reagents as to avoid cross-contaminations. Never pipette solutions by mouth.

10.	The use of automatic pipettes is recommended.

11.	Assure that the incubation temperature is 37°C inside the incubator.

12.	When adding samples, avoid touching the well’s bottom with the pipette tip.

13.	When reading the results with a plate reader, it is recommended to determine the absorbance at 450nm or at 450nm with reference at 630nm.

14.	All specimens from human origin should be considered as potentially infectious.

15.	Materials from human origin may have been used in the kit. These materials have been tested with tests kits with accepted performance and found negative for antibodies to HIV 1/2, HCV, TP and HBsAg. However, there is no analytical method that can assure that infectious agents in the specimens or reagents are completely absent. Therefore, handle reagents and specimens with extreme caution as if capable of transmitting infectious diseases. Strict adherence to GLP (Good Laboratory Practice) regulations can ensure the personal safety. Never eat, drink, smoke, or apply cosmetics in the assay laboratory.

16.	Bovine derived sera may have been used in this kit. Bovine serum albumin (BSA) and fetal calf sera (FCS) are derived from animals from BSE/TSE free-geographical areas.

17.	The pipette tips, vials, strips and sample containers should be collected and autoclaved for 1 hour at 121°C or treated with 10% sodium hypochlorite for 30minutes to decontaminate before any further steps for disposal.

18.	The Stop solution (2M H2SO4 ) is a strong acid. Corrosive. Use it with appropriate care. Wipe up spills immediately or wash with water if come into contact with the skin or eyes. ProClin 300 used as a preservative can cause sensation of the skin.

19.	The enzymatic activity of the HRP-conjugate might be affected from dust, reactive chemical, and substances like sodium hypochlorite, acids, alkalins etc. Do not perform the assay in the presence of such substances.

20.	Materials Safety Data Sheet (MSDS) available upon request.

21.	If using fully automated microplate processing system, during incubation, do not cover the plates with the plate cover. The tapping out of the remainders inside the plate after washing, can also be omitted.

ASSAY PROCEDURE

Step1 Reagents preparation: Allow the reagents to reach room temperature (18-30°C). Check the Wash buffer concentrate for the presence of salt crystals. If crystals have formed in the solution, resolubilize by warming at 37`C until crystals dissolve. Dilute the stock Wash Buffer 1 to 20 with distilled or deionized water. Use only clean vessels to dilute the buffer.

Step2	Numbering Wells: Set the strips needed in

strip-holder and number sufficient number of wells including three Negative control (e.g. B1, C1, D1), two Positive control (e.g. E1, F1) and one Blank (e.g. A1, neither samples nor HRP-Conjugate should be added into the Blank well). Use only number of strips required for the test.

Step3 Adding Sample: Add 5011l of Positive control, Negative control, and samples into their respective wells. Note: Use a separate disposal pipette tip for each specimen, Negative Control and Positive Control to avoid cross-contamination.

Step4 Adding HRP-Conjugate:Add 5011l HRP-Conjugate to each well except the Blank and mix by tapping the plate gently.

Step5 Incubating: Cover the plate with the plate cover and incubate for 30 minutes at 37°C.It is recommended to use water tank to assure the temperature stability and humidity during the incubation. If dry incubator is used, do not open the door frequently.

Step6 Washing: At the end of the incubation, remove and discard the plate cover. Wash each well 5times with diluted Wash buffer. Each time allow the microwells to soak for 30-60 seconds. After the final washing cycle, turn the strips plate down onto blotting paper or clean towel, and tap the plate to remove any remainders.

Step7 Coloring: Dispense 5011l of Chromogen A and 5011l Chromogen B solution into each well including the Blank, and mix by tapping the plate gently. Incubate the plate at 37°C for 30minutes avoiding light. The enzymatic reaction between the Chromogen solutions and the HRP-Conjugate produces blue color in Positive control and Positive sample wells.

Step8 Stopping Reaction: Using a multichannel pipette or manually add 5011l Stop Solution into each well and mix gently. Intensive yellow color develops in Positive control and Positive sample wells.

Step9 Measuring the Absorbance: Calibrate the plate reader with the Blank well and read the absorbance at 450nm.If a dual filter instrument is used, set the reference wavelength at 630nm. Calculate the Cut-off value and evaluate the results. (Note: read the absorbance within 10 minutes after stopping the reaction)

INTERPRETATION OF RESULTS AND QUALITY CONTROL

Each microplate should be considered separately when calculating and interpreting results of the assay, regardless of the number of plates concurrently processed. The results are calculated by relating each sample optical density (OD) value to the Cut-off value (C.O.) of the plate. If the Cut-off reading is based on single filter plate reader, the results should be calculated by subtracting the Blank well OD value from the print report values of samples and controls. In case the reading is based on Dual filter plate reader, do not subtract the Blank well OD from the print report values of samples and controls.

1. Calculation of Cut-off value:

Cut-off value (C.O.) = *Nc x 0.3

*Nc = the mean absorbance value for three negative controls.

Important: If the mean OD value of the negative control is higher than 2.500, take it as 2.500 .If lower than 2.500, see the Quality control range.

Example:

Example:

| 1. Calculation of Nc: |  |  |  |
| --- | --- | --- | --- |
| Well No | B1 | C1 | D1 |
| Negative controls OD value | 1.482 | 1.437 | 1.479 |

| 1. Calculation of Nc: |  |  |  |
| --- | --- | --- | --- |
| Well No | B1 | C1 | D1 |
| Negative controls OD value | 1.482 | 1.437 | 1.479 |

Nc=1.466

2.	Calculation of Cut-off value: (C.O.)= 1.466/2= 0.733

Nc=1.466

2.	Calculation of Cut-off value: (C.O.)= 1.466/2= 0.733

If one of the Negative control values does not meet the Quality control range specifications, it should be discarded and the mean value is calculated again using the remaining two values. If more than one control OD value does not meet the Quality control range specifications, the test is invalid and must be repeated.

2. Quality control range:

The test results are valid if the Quality Control criteria are verified. It is recommended that each laboratory must establish appropriate quality control system with quality control material similar to or identical with the patient sample being analyzed

1.	The absorbance of the Blank well, which contains only Chromogens and Stop solution, is less than 0.080 at 450 nm.

2.	The absorbance value OD of the Positive control must be lower than 0.100 at 450/630nm or at 450nm after blanking.

3.	The absorbance value OD of the Negative control must equal to or higher than 1.000 at 450/630nm or at 450nm after blanking.

3. Interpretations of the results:

(S = the individual absorbance (OD) of each specimen)

Negative Results	(S/C.O. ≥ 1): samples giving an

absorbance higher than or equal to the Cut-off value are
considered negative, which indicates that antibodies to the
hemagglutinin of H5 avian influenza have not been detected

with this kit.

Positive Results (S/C.O. < 1): samples giving an absorbance less than the Cut-off value are considered initially reactive, which indicates that antibodies to hemagglutinin of H5 avian influenza have probably been detected with this kit. Any initially reactive samples must be retested in duplicates. Repeatedly reactive samples can be considered positive for H5 HA specific antibodies.

Borderline: Samples with absorbance to Cut-off ratio between 0.9 and 1.00 are considered borderline samples and retesting is recommended. Repeatedly positive samples can be considered positive for H5 HA antibodies.

LIMITATIONS

1.	Non-repeatable positive result may occur due to the general biological and biochemical characteristics of ELISA assays. The test is designed to achieve performance characteristics of high sensitivity and specificity. HA may be undetectable during the early stages of the disease and in some immunosuppressed individuals.

2.	Any positive results must be interpreted in conjunction with patient clinical information and other laboratory testing results.

3.	Common sources for mistakes: kits beyond the expiry date, bad washing procedures, contaminated reagents, incorrect assay procedure steps, insufficient aspiration during washing, failure to add samples or reagents, equipment, timing, volumes, sample nature and quality.

4.	If, after retesting of the initially reactive samples, the assay results are negative , these samples should be considered as non-repeatable (false positive) and interpreted as negative. As with many very sensitive ELISA assays, false positive results can occur due to the several reasons, most of which are related but not limited to inadequate washing step. For more information regarding ELISA Troubleshooting, please refer to ’s “ELISAs and Troubleshooting Guide”, or contact Beijing technical support for further assistance.

5.	The prevalence of the marker will affect the assay’s predictive values.

6.	This kit is intended ONLY for testing of individual serum or plasma samples. Do not use it for testing of cadaver samples, saliva, urine or other body fluids, or pooled (mixed) blood.

VALIDITY

Please do not use this kit beyond the expiry date indicated on the kit box and reagent labels.

INDICATIONS OF INSTABILITY OR DETERIORATION OF
THE REAGENTS

1.	Values of the Positive or Negative controls ,which are out of the indicated Quality control range, are indicator of possible deterioration of the reagents and/or operator or equipment errors. In such case, the results should be considered as invalid and the samples must be retested. In case of constant erroneous results classified as due to deterioration or instability of the reagents, immediately substitute the reagents with new ones, or contact Beijing technical support for further assistance.

2.	If after mixing of the Chromogen A and B solutions into the wells, the, the color of the mixture turns blue within few

minutes, do not continue carrying out the testing and replace the reagents with fresh ones.

REFERENCES:

1.	Fouchier RAM, Schneeberger PM, Rozendaal FW, et al. Avian influenza A virus (H7N7) associated with human conjunctivitis and a fatal case of acute respiratory distress syndrome. Proc Natl Acad Sci 2004. Published online before print January 26, 2004

2.	Fouchier RAM, Munster V, Wallensten A, et al. Characterization of a novel influenza A virus hemagglutinin subtype (H16) obtained from black-headed gulls J Virol 2005 Mar;79(5):2814-22

3.	Horimoto T, Kawaoka Y. Pandemic threat posed by avian influenza A viruses. Clin Microbiol Rev 2001;14(1):129-49

4.	Keawcharoen J, Oraveerakul K, Kuiken T, et al. Avian influenza H5N1 in tigers and leopards. Emerg Infect Dis 2004 Dec;10(12)

5.	Luschow D, Werner O, Mettenleiter TC, et al. Protections of chickens from lethal avian influenza A virus infections by live-virus hemagglutinin (H5) gene. Vaccine 2001 Jul 20;19(30):4249-59

6.	Monto AS. The threat of an avian influenza pandemic. (Perspective) N Engl J Med 2005 Jan 27;352(4):323-4

7.	OIE. Highly pathogenic avian influenza. Technical disease card database. Apr 22, 2002

8.	OIE. Highly pathogenic avian influenza. International Health Code. Chap 2.7.12. 2004

9.	OIE. Update on avian influenza in animals in Asia (type H5). Updated frequently

10.	Subbarao K, Klimov A, Katz J, Regnery H, Lim W, Hall H, et al. Characterization of an avian influenza A (H5N1) virus isolated from a child with a fatal respiratory illness. Science 1998;279:393–6.

11.	Yuen KY, Chan PKS, Peiris M, Tsang DNC, Que TL, Shortridge KF, et al. Clinical features and rapid viral diagnosis of human disease associated with avian influenza A H5N1 virus. Lancet 1998;351:467–71.

FOR RESEARCH USE ONLY

| SUMMARY OF THE ASSAY PROCEDURE: |
| --- |
| Add Sample | 50µl |
| Add HRP-Conjugate | 50µl |
| Incubate 37°C | 30minutes |
| Wash | 5times |
| Coloring | 50µl A + 50µl B |
| Incubate 37°C | 30minutes |
| Stop the reaction | 50µl stop solution |
| Read the absorbance | 450nm or 450/630nm |

| SUMMARY OF THE MAJOR COMPONENTS OF THE KIT: |
| --- |
| Microwell plate Negative/ Positive control HRP-Conjugate Wash Buffer Chromogen A/B/ Stop Solution | One/ 96 wells One each/ 0.5ml One/ 6ml One/ 50ml One each/ 6ml |
| Note: the components of individual kits are not interchangeable |