Competitive Enzyme Immunoassay Kit for

Quantitative Analysis of Salinomycin

1. Background

Salinomycin is commonly used as anti-coccidiosis in chicken.It leads to vasodilatation, especially coronary artery expansion and blood flow increment, which has no side effects on normal people, but for those who have got coronary artery diseases, it can be very dangerous to eat food with too much Salinomycin, the result is vasodilatation, partial absence of oxygen in heart and health deterioration. So Salinomycin has been an important drug monitored an controlled in drug residue analysis of animal derived food analysis and feed.

Salinomycin distributes widely inside animal body, especially in liver and fat, the concentration is very high, the secondary is muscle, kidney and plasma. HPLC is the common method of detecting Salinomycin, and it costs too much compared to ELISA, which is more sensitive and precise and more appropriate for on-site supervision and mass screening.

This kit is a new product for drug residual detection based on ELISA technology, which is fast, easy to process, precise and sensitive, and it can considerably minimize operation errors and work intensity

2. Test Principle

This kit is based on indirect-competitive ELISA technology. The microtiter wells are coated with coupling antigen. Salinomycin residue in the sample competes with the antigen coated on the microtiter plate for the antibody. After the addition of enzyme labeled anti-antibody, TMB substrate is used to show the color. Absorbance of the sample is negatively related to the Salinomycin reside in it, after comparing with the Standard Curve, multiplied by the dilution multiple, Salinomycin residue quantity in the sample can be calculated.

3. Applications

This kit can be used in quantitative and qualitative analysis of Salinomycin residue in animal tissue(muscle, liver).

4. Cross-reactions

Salinomycin…………………………………100%

Narasin………………………………80%

maduramicin…………………………< 1%

erythromycin …………………………< 1%

tylosin…………………………………< 0.1%

monensin……………………………< 0.1%

5. Materials Required

5.1 Equipments:

┅┅Microtiter plate spectrophotometer (450nm/630nm)

┅┅Rotary evaporator or nitrogen drying instruments

┅┅Homogenizer or stamocher

┅┅Shaker

┅┅Centrifuge

┅┅Micropipettes: 20(l～200(l、100(l～1000(l

250(l-multipipette

┅┅Analytical balance (inductance: 0.01g)

┅┅Glass test tube: 10ml

┅┅Graduated pipette: 10ml

┅┅Polystyrene Centrifuge tubes:  50ml

5.2 Reagents:

┅┅carbinol(AR)

┅┅methylene chloride (AR)

┅┅sodium hydrate (AR)

┅┅sodium chloride(AR)

┅┅deionized water

6. Kit Components

1、Microtiter plate with 96 wells coated with coupling antigen

2、standard solutions(6 bottles×1ml/bottle)

0ppb, 1ppb, 3ppb, 9ppb, 27ppb, 81ppb

3、 High concentration standard control:(1ml/瓶)  1ppm

4、Enzyme-labeled secondary antibody solution 12ml…red cap

5、 antibody work  solution   7ml………… green cap

6、Solution A 7ml…… white cap

7、Solution B 7ml …… red cap

8、Stop solution 7ml ………… yellow cap

9、 20×concentrated wash solution 40ml ………transparent cap

10、2×concentrated extraction solution 50ml…… blue cap

7. Reagents Preparation:

Solution 1: 8% sodium chloride solution

Weigh 8 g sodium chloride and dilute to 100ml with deionized water.

Solution 2: carbinol－8% sodium chloride solution

Add 80ml carbinol to 20ml  8% sodium chloride solution, and mix completely.

Solution 3: 0.5 M sodium hydrate－8%sodium chloride solution

Dissolve 2 g sodium hydrate with 8% sodium chloride solution to 100ml.

Solution 4: extraction solution

Dilute the 2×concentrated extraction solution with deionized water in the volume ration of 1:1, which will be used for sample extraction. the diluted extraction solution can be stored at 4℃ for 1 month.

Solution 5: wash solution

Dilute the 20×concentrated wash solution with deionized water in the volume of 1:19 according to your requirement. It will be used for washing the plates.The diluted wash solution can be stored at 4℃ for 1 month.

8. Sample Preparations

8.1 Notice and precautions for the users before operation:

(a) Please use one-off tips in the process of experiment, and change the tips when absorb different reagent.

(b) Make sure that all experimental instruments are clean.

(c) Keep untreated samples in freeze;

(d) Treated samples should be tested immediately ;

8.2 Tissue samples( muscle and liver):

----Take some of the sample and homogenate at 4000r/min for 1min with homogenizer or stamocher;

----Weigh 2.0±0.05g homogenate into a 50ml polystyrene centrifuge tube, add 8ml carbinol－8%sodium chloride(See solution 2),shake for 5min with shaker.

----Centrifuge at least at 3000g, at room temperature(20-25℃) for 10min;

----Transfer 2ml of the supernate and mix with 4ml 0.5 M sodium hydrate —8% sodium chloride(See solution 3), then add 10ml methylene chloride, shake for 5min with shaker, centrifuge at least 3000g, at room temperature (20-25℃) for 10min.

----Remove the supernatant liquid, transfer 5ml of the substrate into a 10ml clean glass tube, dry with 50～60℃ water bath nitrogen gas flow.

----Dissolve the dry leftover with 1ml extraction solution(See solution 4).

----Take 50(l of the prepared solution for assay.

8.3 eggs

┅┅Homogenize the egg white with yolk;

┅┅Transfer 1ml of the homogenate into a 10ml graduated centrifuge pipette , add 5ml acetonitrile, shake for 5min with shaker, centrifuge at least 3000g, at room temperature for 10min.

┅┅Transfer 1ml of the supernate into a 10ml clean glass tube, dry with 50～60℃ water bath nitrogen gas flow.

┅┅Dissolve the dry leftover with  1ml extraction solution (See solution 4),then whorl for 30s with gyroscope  or vortex mixer.

┅┅Mix 200μl of the dissolved solution with 600μl extraction solution, mix completely.

┅┅Take 50μl of the prepared solution for assay.

9. Assay steps:

9.1 Notice before assay:

9.1.1 Make sure all reagents and microwells are all at room temperature (20-25℃).

9.1.2 Return all the rest reagents to 2～8℃ immediately after used.

9.1.3 Washing the microwells correctly is an important step in the process of assay; it is the vital factor to the reproducibility of the ELISA analysis.

9.1.4 Avoid the light and cover the microwells during incubation.

9.2 Assay Steps:

9.2.1 Take all reagents out at room temperature (20-25℃) for more than 30min, homogenize before use.

9.2.2 Get the microwells needed out and return the rest into the zip-lock bag at 2-8℃ immediately.

9.2.3 The diluted wash solution should be rewarmed to room temperature before use.

9.2.4 Number: Number every microwell position and all standards and samples should be run in duplicate. Record the standards and samples positions.

9.2.5 Add standard solution/sample: Add 50 µl of standard solution or prepared sample to corresponding wells. Add 50µl antibody work solution. Mix gently by shaking the plate manually and incubate for 30min at 25℃ with cover.

9.2.6 Wash: Remove the cover gently and pure the liquid out of the wells and rinse the microwells with 250µl diluted wash solution (solution 5) at interval of 10s for 4~5 times. Absorb the residual water with absorbent paper (the rest air bubble can be eliminated with unused tip).

9.2.7 Enzyme-labeled secondary antibody solution: Add enzyme-labeled secondary antibody solution 100(l to each well, Mix gently by shaking the plate manually and incubate for 30min at 25℃ with cover. Repeat the wash step again.

9.2.8 Coloration: Add 50µl solution A and 50µl solution B to each well. Mix gently by shaking the plate manually and incubate for 15 min at 25℃ with cover(see 12.8).

9.2.9 Measure: Add 50µl the stop solution to each well. Mix gently by shaking the plate manually and measure the absorbance at 450nm against an air blank (It’s suggested measure with the dual-wavelength of 450/630nm. Read the result within 5min after addition of stop solution. ) (We can also measure by sight without stop solution in short of the ELIASA instrument)

10. Results

There are 2 different methods to determinate the results. Method 1 leads to a round estimation, and method 2 leads to definite quantity estimation. (Please notice: the absorption is inversely proportional to the Salinomycin concentration in the sample.)

10.1 Round estimation

We can get the range of different strengths from the compare of average absorption and standards by sight. For example, the absorption of sample 1 is 0.611, sample 2 is 1.012, and the absorptions of Salinomycin standard solutions: 1.889(0ppb); 1.391(1ppb);0.947(3ppb);0.483(9ppb);0.225(27ppb); 0.080 (81ppb).  So we can say the strength of diluted sample 1 is between 3ppb and 9ppb; and diluted sample 2 between 1 ppb to 3ppb. In order to obtain the Salinomycin actually contained in a sample, the diluted sample results must be further multiplied by the corresponding dilution multiple.

10.2 Definite quantity estimation

(1) The mean values of the absorbance values obtained for the standards and the samples are divided by the absorbance value of the first standard (zero standard ) and multiplied by 100%. The zero standard is thus made equal to 100% and the absorbance values are quoted in percentages.

B

Absorbance (%) = —— ×100%

B0

B ——absorbance standard (or sample)

B0 ——absorbance zero standard

(2) To draw a standard curve: Take the absorbance value of standards as y-axis, semi logarithmic of the concentration of the Salinomycin standards solution (ppb) as x-axis.

--- The Salinomycin concentration of each sample (ppb), which can be read from the calibration curve, is multiplied by the corresponding dilution multiple of each sample followed, and the actual concentration of sample is obtained.

For evaluation of the ELISA kits special software has been developed for exact and rapid analysis. The analysis software can be ordered on request.

Dilution multiple of samples：

Tissue sample( muscle and liver)：5

Egg sample：20

11. Sensitivity, accuracy and precision

Test Sensitivity : 1ppb

Detection limit:

Tissue sample( muscle and liver) ……5ppb

eggs………………………………………20ppb

Accuracy:

Chicken and chicken liver…………………75±10%

Pork and pig liver…………………………75±10%

eggs…………………………………80±20%

Precision:

Variation coefficient of the ELISA kit is less than 10%.

12. Notice

12.1 The mean values of the absorbance values obtained for the standards and the samples will be reduced if the reagents and samples have not been regulated to room temperature (20-25℃).

12.2 Do not allow microwells to dry between steps to avoid unsuccessful repetitiveness and operate the next step immediately after tap the microwells holder.

12.3 homogenize each reagents before use.

12.4 Keep your skin away from the stop solution for it is the 2M H2SO4 ;

12.5 Don’t use the kits out of date. Don’t exchange the reagents of different batches, or else it will drop the sensitivity.

12.6 Storage condition:

Keep the ELISA kits at 2-8℃,do not freeze. Seal rest microwell plates Avoid straight sunlight for the standard sample and the colorless chromogenic reagent are sensitive to light.

12.7 Indications for the reagents go bad:

Substrate solution should be abandoned if it turns colors.

The reagents may be turn bad if the absorbance value (450/630nm) of the zero standard is less than 0.5(A450nm＜0.5).

12.8 The coloration reaction needs 10-15min after adding Solution A and Solution B. And you can prolong the incubation time ranges from 10-15min to 20min or more if the color is too light to be determined., never exceed 30min,On the contrary, shorten the incubation time properly.

12.9 The optimal reaction temperature is 25℃. Higher or lower temperature will lead to the changes of sensitivity and absorbance values.

13. Storage condition and storage period

Storage condition: 2-8℃.

Storage period: 6 months.