Competitive Enzyme Immunoassay Kit for

Quantitative Analysis of Nitrofurazone metabolite (SEM)

1. Background

Nitrofurans are very good antibacterial, which have been broadly applied in poultry, aquatic and swine production. Long term research indicates that Nitrofurans and their metabolites lead to caner and gene mutations in lab animals, thus these drugs are being prohibited in therapy and feedstuff.

Nitrofurans are metabolized inside the body very fast, and their metabolites combined with the tissues would exist for quite a long time, so residual analysis of these drugs will depend on the detection of their metabolites, including Furazolidone metabolite (AOZ), Furaltadone metabolite (AMOZ), Nitrofurantoin metabolite (AHD) and Nitrofurazone metabolite (SEM).

The common approach to detect nitrofurazone metabolite is LC-UV, LC-MS and LC-MS/MS. The ELISA test, in which specific antibody of SEM ramification is used is more accurate, sensitive, and simple to operate.

2. Test Principle

This ELISA kit is designed to detect nitrofurazone metabolite (SEM) based on the principle of “indirect-competitive” enzyme immunoassay. The microtiter wells are coated with capture BSA-linked antigen. SEM in sample competes with the antigen coated on the microtiter plate for the antibody. After the addition of horseradish peroxidase labeled anti-antibody, tetramethyl-benzidine (TMB)/peroxide substrate is used and the signal is measured by spectrophotometer. The absorption is inversely proportional to the SEM concentration in the sample.

3. Applications

This kit can be used for the quantitative and qualitative analysis of SEM in aquatic products.

4. Cross-reactions

Nitrofurazone metabolite (SEM)………………100%

nitrofurazone……………………………………25%

furazolidone metabolite………………………<0.1%

Furaltadone metabolite………………………<0.1%

Nitrofurantoin metabolite………………………<0.1%

furazolidone……………………………………………<1%

Furaltadone………………………………………………<1%

Nitrofurantoin……………………………………………<1%

5. Materials Required

5.1 Equipments

---Microtiter plate spectrophotometer (450nm/630nm)

---Rotary evaporator or nitrogen drying instruments

---Homogenizer (or stomacher)

---Shaker

---Gyroscope (or vortex mixer)

---Centrifuge

---Analytical balance (inductance: 0.01g)

---Graduated pipette: 10ml

---Rubber pipette bulb

---Volumetric flask: 100ml, 500ml

---Glass test tube: 10ml

---Polystyrene centrifuge tube: 50ml

---Micropipettes: 20ul-200ul, 100ul-1000ul,

250ul multi-channel pipette

5.2 Reagents

---Ethyl acetate (AR)

---n-Hexane (or normal heptane) (AR)

---Potassium phosphate dibasic trihydrate (K2HPO4.3H2O ) (AR)

---methanol (AR)

---Concentrated hydrochloric acid

---Sodium hydroxide (NaOH) (AR)

---Deionized water

6. Kit Components

1. Microtiter plate with 96 wells coated with coupling antigen

2.  SEM standard solutions×6 bottles: (1ml/bottle)

0ppb, 0.05ppb, 0.15ppb, 0.45ppb, 1.35ppb, 4.05ppb

3. High concentration standard control: (1ml/bottle) 100ppb

4. Enzyme conjugate 12ml ………………………………red cap

5. Antibody solution 7ml …………………………green cap

6. Solution A 7ml ……………………………white cap

7. Solution B 7ml ……………………………………red cap

8. Stop solution  7ml …………………………yellow cap

9.20×Concentrated wash solution 40ml……transparent cap

10. 2×Concentrated extraction solution   50ml……blue cap

11. 2- nitrobenzaldehyde    15.1mg……black cap

7. Reagents Preparation

Solution 1: Derivative Reagent: Add methanol to the bottle with 2-nitrobenzaldehyde and dilute to 10ml. (the final concentration is 10mM)

Solution 2: 0.1M Potassium phosphate dibasic: 22.8g p otassium phosphate dibasic trihydrate (K2HPO4.3H2O ) dissolved with deionized water, dilute to 1L.

Solution 3: 1M Hydrochloric acid solution:

Dilute the concentrated hydrochloric acid 8.3ml with deionized water and dilute to 100ml.

Solution 4: 1M Sodium hydroxide (NaOH):

4g sodium hydroxide (NaOH) dissolved with deionized water and dilute to 100ml.

Solution 5: methanol solution

Mix 90ml methanol with 10ml deionized water;

Solution 6: extraction solution

Dilute the concentrated extraction solution with deionized water in the volume ratio of 1:1, which will be used for sample extraction. The diluted extraction solution can be conserved for a month at 4℃.

Solution 7: Wash solution

Dilute the 20×concentrated wash solution with deionized water in the volume ratio of 1:19, which will be used to rinse the plate. This solution can be conserved at 4℃ for 1 month;

8. Sample Preparation

8.1 Notice and precautions for the users before operation:

(a) Please use one-off tips in the process of experiment, and change the tips when absorbing different reagent.

(b) Make sure that all experimental instruments are clean.

(c) The derivative reagent can be conserved at 2-8℃ for half a year;

(d) Keep untreated samples in freeze;

(e) Treated samples can be conserved for 24h at 2-8℃ in darkness.

(f) The diluted derivative reagent can be conserved at 4℃ for half a year;

(g) Potassium phosphate dibasic solution can be conserved at 2-8℃ for three months;

(h) The hydrochloric acid solution can be conserved at room temperature for 3 months;

(i) The sodium hydroxide solution can be conserved for 3 months at room temperature;

8.2Aquatic products

┅┅Homogenize samples with homogenizer.

┅┅Weigh 1.0 ±0.05g of the homogenate

┅┅Add 5ml methanol solution(see solution 5),shake for

5min, then centrifuge for 10min, at least 3000g, remove all the supernate (this step can decrease the ground substance interference);

┅┅Add 4ml deionized water, whorl with gyroscope to dissolve completely, add 0.5ml 1M hydrochloric acid solution (see solution 3) and 100(l derivative reagent (see solution 1), and shake completely with shaker;

┅┅Incubate at 37℃ over night (about 16 hours)

┅┅Then add 5ml 0.1M Potassium phosphate dibasic solution (see solution 2), 0.4ml 1M sodium hydroxide solution (see solution 4) and 10ml ethyl acetate, shake fiercely for 30s with shaker;

┅┅Centrifuge at room temperature(20-25℃) for 10 min, at least 3000g;

┅┅Take 5ml of the ethyl acetate phase into a 10ml clean

glass test tube, dry with 50-60 ℃ nitrogen gas flow;

┅┅Dissolve the dry leftover with 1ml n-hexane(or normal heptane),whorl for 30s with gyroscope , then add 0.5ml extraction solution (see solution 6) , whorl for 1min to mix completely┅┅Centrifuge at room temperature(20-25℃) for 10min, at least 3000g;

┅┅Remove the supermatant organic phase ans take 50µl of the substrate water phase for assay;

9. Assay process

9.1 Notice before assay:

9.1.1 Make sure all reagents and microwells are all at room temperature ( 20-25℃)

9.1.2 Return all the rest reagents to 2-8℃immediately after used.

9.1.3 Washing the microwells correctly is an important step in the process of assay; it is the vital factor to the reproducibility of the ELISA analysis.

9.1.4 Avoid the light and cover the microwells during incubation.

9.2 Assay Steps:

9.2.1 Take all reagents out at room temperature ( 20-25℃)for more than 30min, homogenize before use.

9.2.2 Get the microwells needed out and return the rest into the zip-lock bag at 2-8℃ immediately.

9.2.3 The diluted wash solution should be rewarmed to be at room temperature before use.

9.2.4 Number: Numbered every microwell positions and all standards and samples should be run in duplicate. Record the standards and samples positions.

9.2.5 Add standard solution or sample solution:  add standard solution or sample solution 50(l to the corresponding well, then add antibody solution 50(l to each well, mix gently, incubate for 30min at 4℃ in darkness with cover ;

9.2.6 Wash: Remove the cover gently and pour the liquid out of the wells and rinse the microwells with 250µl diluted wash solution (see solution 7) at interval of 50s –1minfor 4~5 times. Absorb the residual water with absorbent paper.

9.2.7 Add enzyme labeled secondary antibody: add enzyme labeled antibody solution 100(l to each well, mix gently,then incubate for 30 min at 25℃ in darkess with cover , take out and repeat the wash step;

9.2.8 Coloration: add 50µl solution A and 50µl solution B to each well. Mix gently by shaking the plate manually and incubate for 15min at 25℃ with cover (see 12.8).

9.2.9 Measure: add 50µl of the stop solution to each well. Mix gently by shaking the plate manually and measure the absorbance at 450nm against an air blank (It’s suggested measure with the dual-wavelength of 450/630nm. Read the result within 5min after addition of stop solution. ) (We can also measure by sight without stop solution in short of the ELIASA instrument)

10. Results

There are 2 different methods to determinate the results. Method 1 leads to a round estimation, and method 2 leads to definite quantity estimation. (Please notice: the absorption is inversely proportional to the SEM concentration in the sample.)

10.1 Round estimation

We can get the range of different strengths from the compare of average absorption and standards by sight. For example, the absorption of sample 1 is 0.236, sample 2 is 0.666, and the absorptions of SEM standard solutions: 1.949(0ppb); 1.434(0.05ppb); 0.939(0.15ppb); 0.487(0.45ppb); 0.157(1.35ppb); 0.060(4.05ppb). So we can say the strength of diluted sample 1 is between 0.45ppb and 1.35ppb; and diluted sample 2 between 0.15ppb to 0.45ppb. In order to obtain the SEM actually contained in a sample, the diluted sample results must be further multiplied by the corresponding dilution factor.

10.2 Definite quantity estimation

(1) The mean values of the absorbance values obtained for the standards and the samples are divided by the absorbance value of the first standard (zero standard ) and then multiplied by 100%. The zero standard is thus made equal to 100% and the absorbance values are quoted in percentages.

B ——absorbance standard (or sample)

B0 ——absorbance zero standard

(2) To draw a standard curve: Take the absorbance value of standards as y-axis, semi logarithmic of the concentration of the SEM standards solution (ppb) as x-axis.

--- The SEM concentration of each sample (ppb), which can be read from the calibration curve, is multiplied by the corresponding dilution factor of each sample followed, and the actual concentration of sample is obtained.

Please notice:

For evaluation of the ELISA kits, special software has been developed for exact and rapid analysis. The analysis software can be ordered on request.

Dilution factor of samples:

Fish, shrimp, liver and muscle samples: 1

11. Sensitivity, accuracy and precision

Test Sensitivity: 0.05ppb

Detection limit:

Tissue sample…0.1ppb

Tissue samples of fish and shrimp have some interference, so the detection limit is 0.2 ppb;

Accuracy:

Fish, shrimp, liver and muscle samples………… 80±15%

Precision:

Variation coefficient of the ELISA kit is less than 10%.

12. Notice

12.1 The mean values of the absorbance values obtained for the standards and the samples will be reduced if the reagents and samples have not been regulated to room temperature  ( 20-25℃)

12.2 Do not allow microwells to dry between steps to avoid unsuccessful reproducibility and operate the next step immediately after tap the microwells holder.

12.3. Homogenize each reagent before using.

12.4. Keep your skin away from the stop solution for it is the 2M H2SO4 solution.

12.5 Don’t use the kits out of date. Don’t exchange the reagents of different batches, or else it will drop the sensitivity.

12.6 Storage condition:

Keep the ELISA kits at ( 2-8℃), do not freeze. Seal rest microwell plates, Avoid straight sunlight during all incubations. Covering the microtiter plates is recommended.

12.7 Indications for the reagents going bad:

Substrate solution should be abandoned if it turns colors.

The reagents may be turn bad if the absorbance value (450/630nm) of the zero standard is less than 0.5(A450nm＜0.5).

12.8 The coloration reaction needs 15min after the addition of solution A and solution B; but you can prolong the incubation time to 20min or more if the color is too light to be determined. never exceed 30min,On the contrary, shorten the incubation time properly.

12.9 The optimal reaction temperature of adding SEM standard solution, antibody solution is 4℃.  And after adding enzyme labeled secondary antibody solution and Solution A, Solution B, the optimal reaction temperature is 25℃. Higher or lower temperature will lead to the changes of sensitivity and absorbance values.

13. Storage

Storage condition: 2-8.℃

Storage period:  6 months.