Competitive enzyme immunoassay Kit for

Quantitative Analysis of Thymopetide

1. Background

Thymopetide is a series of biologicalactive peptides that excreted by thymus, it is used for immunity regulation.

this kit is a new product for drug residual detection based on ELISA technology, which only costs 1.5 hours in each operation and can considerably minimize operation errors and work intensity, and can be used for determination of  Thymopetide in serum.

2. Test Principle

This kit is based on indirect-competitive ELISA technology. The microtiter wells are coated with coupling antigen. Thymopetide residual in the sample competes with the antigen coated on the microtiter plate for the antibody. After the addition of horseradish peroxidase labeled anti-antibody, tetramethyl-benzidine (TMB) substrate is used to show the color. Absorbance of the sample is negatively related to the Thymopetide residual in it, after comparing with the Standard Curve, multiplied by the dilution multiples, Thymopetide residual quantity in the sample can be calculated.

3. Applications

This kit can be used in quantitative and qualitative analysis of Thymopetide residue in serum and ascites.

4. Cross-reactions

Thymopetide………………………………………100%

thymopentin…………………………………………<1%

interferon……………………………………………<1%

5. Materials Required

5.1 Equipments:

┅┅Microtiter plate spectrophotometer (450nm/630nm)

┅┅Graduated pipette: 10ml

┅┅Rubber pipette bulb

----Polystyrene centrifuge tube: 2ml

┅┅Micropipettes: 20(l～200(l、100(l～1000(l

250(l-multipipette

Reagents：

----Sodium hydrogen phosphate dodecahydate (AR)

---- Sodium dihydrogen phosphate dihydrate (AR)

┅┅deionized water

6. Kit Components

1、Microtiter plate with 96 wells coated with coupling antigen

2、standard solutions(6 bottles:1ml/bottle)

0ppb, 1ppb, 4 ppb, 16 ppb, 64 ppb, 256 ppb

3、High concentration standard control: (1ml/bottle) 1ppm

4、Enzyme labeled secondary antibody solution 12ml… red cap

5、 antibody solution     7ml… green cap

6、 Solution A 7ml… white cap

7、 Solution B 7ml… red cap

8、 Stop solution 7ml… yellow cap

9、 20×Concentrated wash solution   40ml… transparent cap

10、2×Concentrated extraction solution 50ml… blue cap

7. Reagents Preparation:

Solution 1: 0.02M phosphate buffer solution

Dissolve 5.16g Sodium hydrogen phosphate dodecahydate, 0.87g Sodium dihydrogen phosphate dihydrate with deionized water and dilute to 1L;

Solution 2: extraction solution

Dilute the concentrated extraction solution with deionized water in the volume ratio of 1: 1, which will be used for sample extraction. This solution can be conserved for a month at 4℃.

Solution 3: Wash solution

Dilute the concentrated wash solution with deionized water in the volume ratio of 1:19, which will be used to rinse the plates. The diluted wash solution can be conserved for a month at 4℃.

8. Sample Preparations

8.1 Notice and precautions for the users before operation:

(a) Please use one-off tips in the process of experiment, and change the tips when absorbing different reagent.

(b) Make sure that all experimental instruments are clean.

(c) Keep untreated samples in freeze;

(d) Treated samples should be tested immediately.

# 8.2 Serum samples:

┅┅Keep the blood sample at room temperature for 30min, then centrifuge at 10℃ for 10min, at least 3000g, till the serum separates or filter to get the serum;

┅┅Transfer 50(l of the prepared serum into a 2ml clean centrifuge tube then add 1.95ml 0.02 M phosphate buffer solution and mix completely;

┅┅Take 50(l of the prepared solution for assay;

# 9. Assay process

# 9.1 Notice before assay:

9.1.1 Make sure all reagents and microwells are all at room temperature (20-25℃).

9.1.2 Return all the rest reagents to 2～8℃ immediately after used.

9.1.3 Washing the microwells correctly is an important step in the process of assay; it is the vital factor to the reproducibility of the ELISA analysis.

9.1.4 Avoid the light and cover the microwells during incubation.

# 9.2 Assay Steps:

9.2.1 Take all reagents out at room temperature (20-25℃) for more than 30min, homogenize before use.

9.2.2 Get the microwells needed out and return the rest into the zip-lock bag at 2-8℃ immediately.

9.2.3 The diluted wash solution should be rewarmed to be at room temperature before use.

9.2.4 Number: Numbered every microwell positions and all standards and samples should be run in duplicate. Record the standards and samples positions.

9.2.5 Add standard solution/sample: Add 50 µl of standard solution or prepared sample to corresponding wells. Add 50µl antibody solution. Mix gently by rocking the plate manually and incubate for 30min at 37℃ with cover.

9.2.6 Wash: Remove the cover gently and pure the liquid out of the wells and rinse the microwells with 250µl diluted wash solution (solution 4) at interval of 10s for 4~5 times. Absorb the residual water with absorbent paper (the rest air bubble can be eliminated with unused tip).

9.2.7 Enzyme-labeled secondary antibody: Add Enzyme-labeled secondary antibody solution 100(l to each well, mix gently by rocking the plate manually and incubate for 30min at 37℃ with cover. Then repeat the wash step.

9.2.8 Coloration: Add 50µl solution A and 50µl solution B to each well. Mix gently by rocking the plate manually and incubate for 15min at 37℃ with cover (see 12.8).

9.2.9 Measure: Add 50µl of the stop solution to each well. Mix gently by rocking the plate manually and measure the absorbance at 450nm against an air blank (It’s suggested measure with the dual-wavelength of 450/630nm. Read the result within 5min after addition of stop solution. ) (We can also measure by sight without stop solution in short of the ELIASA instrument)

# 10. Results

There are 2 different methods to determinate the results. Method 1 leads to a round estimation, and method 2 leads to definite quantity estimation. (Please notice: the absorption is inversely proportional to the Thymopetide concentration in the sample.)

10.1 Round estimation

We can get the range of different strengths from the compare of average absorption and standards by sight. For example, the absorption of sample 1 is 0.601, sample 2 is 1.002, and the absorptions of Thymopetide standard solutions: 2.043 (0ppb); 1.562(1ppb); 0.997(4ppb); 0.513 (16ppb); 0.246 (64ppb); 0.115(256ppb). So we can say the strength of diluted sample 1 is between 4ppb and 16ppb; and diluted sample 2 between 1ppb to 4ppb. In order to obtain the Thymopetide actually contained in a sample, the diluted sample results must be further multiplied by the corresponding dilution factor.

10.2 Definite quantity estimation

(1)The mean values of the absorbance values obtained for the standards and the samples are divided by the absorbance value of the first standard (zero standard ) and multiplied by 100%. The zero standard is thus made equal to 100% and the absorbance values are quoted in percentages.

B

Absorbance (%) = —— ×100%

B0

B ——absorbance standard (or sample)

B0 ——absorbance zero standard

(2)To draw a standard curve: Take the absorbance value of standards as y-axis, semi logarithmic of the concentration of the Thymopetide standards solution (ppb) as x-axis.

--- The Thymopetide concentration of each sample (ppb), which can be read from the calibration curve, is multiplied by the corresponding dilution factor of each sample followed, and the actual concentration of sample is obtained.

Please notice:

For evaluation of the ELISA kits special software has been developed for exact and rapid analysis. The analysis software can be ordered on request.

Sample dilution factor:40

# 11. Sensitivity, accuracy and precision

Test Sensitivity: 1ppb

Detection limit: 40ppb

# Accuracy:

Serum………………………………………………  70±10%

# Precision:

Variation coefficient of the ELISA kit is less than 10%.

# 12. Notice

12.1 The mean values of the absorbance values obtained for the standards and the samples will be reduced if the reagents and samples have not been regulated to room temperature (20-25℃).

12.2 Do not allow microwells to dry between steps to avoid unsuccessful reproducibility and operate the next step immediately after tap the microwells holder.

12.3. Homogenize each reagent before using.

12.4. Keep your skin away from the stop solution for it is the 2M H2SO4 solution.

12.5 Don’t use the kits out of date. Don’t exchange the reagents of different batches, or else it will drop the sensitivity.

12.6 Storage condition:

Keep the ELISA kits at 2-8℃, do not freeze. Seal rest microwell plates, Avoid straight sunlight during all incubations. Covering the microtiter plates is recommended.

12.7 Indications for the reagents going bad:

Substrate solution should be abandoned if it turns colors.

The reagents may be turn bad if the absorbance value (450/630nm) of the zero standard is less than 0.5(A450nm＜0.5).

12.8 The coloration reaction need 10-15min after the addition of solution A and solution B; But you can prolong the incubation time ranges from 20min to more if the color is too light to be determined., never exceed 30min,On the contrary, shorten the incubation time properly.

12.9 The optimal reaction temperature is 37℃. Higher or lower temperature will lead to the changes of sensitivity and absorbance values.

# 13. Storage condition and storage period

Storage condition: 2-8℃.

Storage period: 6 months.