Competitive Enzyme Immunoassay Kit for

Quantitative Analysis of Monensin

1. Background

Monensin is a polyether antibiotic, which is widely applied in curing animal disease. The drug inside animal body interfere with ionic equilibrium and energy metabolism of cells , which is toxic and will lead to cell denaturalization and putrescence. It is now being restricted in EU,US and China for its residue in animal tissue is harmful to human beings.

This kit is a new product for drug residual detection based on ELISA technology, which only costs 45min in each operation and can considerably minimize operation errors and work intensity.

2. Test Principle

This kit is based on indirect-competitive ELISA technology. The microtiter wells are coated with coupling antigen. Monensin residual in the sample competes with the antigen coated on the microtiter plate for the antibody. After the addition of horseradish peroxidase labeled anti-antibody, tetramethyl-benzidine (TMB) substrate is used to show the color. Absorbance of the sample is negatively related to the Monensin residual in it, after comparing with the Standard Curve, multiplied by the dilution multiples, Monensin residual quantity in the sample can be calculated.

3. Applications

This kit can be used in quantitative and qualitative analysis of Monensin residual in animal tissues (muscle and liver), etc.

4. Cross-reactions

salinomycin   ………………………………………<1%

Maduramicin ………………………………………<1%

5. Materials Required

5.1 Equipments:

┅┅Microtiter plate spectrophotometer (450nm/630nm)

┅┅Rotary evaporator or nitrogen drying instruments

┅┅Homogenizer or stamocher

┅┅Shaker

┅┅Gyroscope  or vortex mixer

┅┅Centrifuge

┅┅Analytical balance (inductance: 0.01g)

┅┅Graduated pipette: 10ml

----volumetric flask: 500ml

----Glass test tube: 10ml

┅┅Polystyrene centrifuge tube: 2ml、50ml

┅┅Micropipettes: 20(l～200(l、100(l～1000(l

250(l-multipipette

5.2 Reagents:

----Sodium chloride (AR)

----Anhydrous acetonitrile (AR)(for chicken , pork samples)

----methanol (AR)(for liver sample of chicken and pig)

----Sodium hydroxide (AR) (for liver sample of chicken and pig)

----methylene chloride (AR) (for liver sample of chicken and pig)

----deionized water

6. Kit Components

1、Microtiter plate with 96 wells coated with coupling antigen

2、standard solutions(6 bottles:1ml/bottle)

0ppb, 1ppb, 3ppb, 9ppb, 27ppb, 81ppb

3、High concentration standard control:  (1ml/bottle) 1ppm

4、 Enzyme labeled secondary antibody solution 7ml red cap

5、 antibody solution   7ml… green cap

6、 Solution A 7ml… white camp

7、 Solution B 7ml… red cap

8、 Stop solution 7ml… yellow cap

9、 20×concentrated wash solution   40ml… transparent cap

10、2×Concentrated extraction solution 50ml… blue cap

7. Reagents Preparation:

Solution 1: acetonitrile solution(for muscle samples)

Mix 80ml anhydrous acetonitrile with 20ml deionized water, mix completely;

Solution 2: 8% Sodium chloride( for liver sample)

Dissolve 40gNaCl with deionized water and dilute to 500ml;

Solution 3: methanol-8% sodium chloride solution (for liver sample)

Mix 80ml methanol with 20ml 8% sodium chloride solution, mix completely;

Solution 4: 0.5M NaOH solution

Dissolve 2g sodium hydroxide with deionized water and dilute to 100ml;

Solution 5: extraction solution

Dilute the concentrated extraction solution with deionized water in the volume ratio of 1: 1, which will be used for sample extraction. This solution can be conserved for a month at 4℃.

Solution 6: Wash solution

Dilute the concentrated wash solution with deionized water in the volume ratio of 1:19(or according to the requirement), which will be used to rinse the plates. The diluted wash solution can be conserved for a month at 4℃.

8. Sample Preparations

8.1 Notice and precautions for the users before operation:

(a) Please use one-off tips in the process of experiment, and change the tips when absorbing different reagent.

(b) Make sure that all experimental instruments are clean.

(c) Keep untreated samples in freeze;

(d) Treated samples should be tested immediately.

8.2 Chicken ,pork samples:

┅┅Homogenize the samples with homogenizer or stamocher;

┅┅Weigh 2.0±0.05g of the homogenate into a 50ml centrifuge tube, add 10ml acetonitrile solution(see solution 1), shake for 5min, centrifuge at room temperature (20-25℃) for 5min , at least 3000g;

┅┅Transfer 1ml of the supernate into a 10ml clean glass test tube, dry with 50～60℃ nitrogen gas flow, then dissolve the dry leftover with 1ml extraction solution(see solution 5),whorl for 30s to dissolve completely;

┅┅Take 200µl of the sample solution, and add 600µl extraction solution(see solution 5), mix completely;

┅┅Take 50µl of the prepared solution for assay ;

8.3 Liver sample( including livers of chicken and pig)

┅┅Homogenize the samples with homogenizer or stamocher ;

┅┅Transfer 1.0±0.05g of the homogenate into a  50ml polystyrene centrifuge tube, add 5ml acetonitrile-8% sodium chloride solution(see solution 3), shake for 5min, centrifuge at room temperature (20-25℃) for 5min , at least 3000g;

┅┅Take 1ml of the supernate , add 2ml 0.5M NaOH solution(see solution 4), mix completely, then add 8ml methylene chloride ,shake for 2min, centrifuge at room temperature (20-25℃) for 5min , at least 3000g;

┅┅Remove the supernatant phase, transfer 2ml of the substrate phase into a 10ml clean glass test tube, dry with 50～60℃ nitrogen gas flow, then add 1ml extraction solution, whorl for 30s;

┅┅Take 50µl of the prepared solution for assay;

9. Assay process

9.1 Notice before assay:

9.1.1 Make sure all reagents and microwells are all at room temperature (20-25℃).

9.1.2 Return all the rest reagents to 2～8℃ immediately after used.

9.1.3 Washing the microwells correctly is an important step in the process of assay; it is the vital factor to the reproducibility of the ELISA analysis.

9.1.4 Avoid the light and cover the microwells during incubation.

9.2 Assay Steps:

9.2.1 Take all reagents out at room temperature (20-25℃) for more than 30min, homogenize before use.

9.2.2 Get the microwells needed out and return the rest into the zip-lock bag at 2-8℃ immediately.

9.2.3 The diluted wash solution should be rewarmed to be at room temperature before use.

9.2.4 Number: Numbered every microwell positions and all standards and samples should be run in duplicate. Record the standards and samples positions.

9.2.5 Add standard solution/sample: Add 50 µl of standard solution or prepared sample to corresponding wells, add enzyme-labeled secondary antibody solution 50µl to each well,  Add 50µl antibody solution. Mix gently by rocking the plate manually and incubate for 30min at 37℃ with cover.

9.2.6 Wash: Remove the cover gently and pure the liquid out of the wells and rinse the microwells with 250µl diluted wash solution (solution 6) at interval of 10s for 4~5 times. Absorb the residual water with absorbent paper (the rest air bubble can be eliminated with unused tip).

9.2.7 Coloration: Add 50µl solution A and 50µl solution B to each well. Mix gently by rocking the plate manually and incubate for 15min at 37℃ with cover(see 12.8).

9.2.8 Measure: Add 50µl of the stop solution to each well. Mix gently by rocking the plate manually and measure the absorbance at 450nm against an air blank (It’s suggested measure with the dual-wavelength of 450/630nm. Read the result within 5min after addition of stop solution. ) (We can also measure by sight without stop solution in short of the ELIASA instrument)

10. Results

There are 2 different methods to determinate the results. Method 1 led to a round estimation, and method 2 lead to definite quantity estimation. (Please notice: the absorption is inversely proportional to the monensin concentration in the sample.)

10.1 Round estimation

We can get the range of different strengths from the compare of average absorption and standards by sight. For example, the absorption of sample 1 is 0.105 sample 2 is 0.679, and the absorptions of monensin standard solutions: 1.513(0ppb); 1.066(1. ppb);0.771(3ppb); 0.358(9ppb); 0.145(27ppb); 0.044(81ppb). So we can say the strength of diluted sample 1 is between  27 ppb and 81ppb; and diluted sample 2 between 3 ppb to 9 ppb. In order to obtain the monensin actually contained in a sample, the diluted sample results must be further multiplied by the corresponding Dilution factor.

10.2 Definite quantity estimation

(1) The mean values of the absorbance values obtained for the standards and the samples are divided by the absorbance value of the first standard (zero standard ) and multiplied by 100%. The zero standard is thus made equal to 100% and the absorbance values are quoted in percentages.

B

Absorbance (%) = —— ×100%

B0

B ——absorbance standard (or sample)

B0 ——absorbance zero standard

(2) To draw a standard curve: Take the absorbance value of standards as y-axis, semi logarithmic of the concentration of the monensin standards solution (ppb) as x-axis.

--- The monensin concentration of each sample (ppb), which can be read from the calibration curve, is multiplied by the corresponding Dilution factor of each sample followed, and the actual concentration of sample is obtained.

Please notice:

For evaluation of the ELISA kits special software has been developed for exact and rapid analysis. The analysis software can be ordered on request.

Dilution multiple of samples：

Chicken, pork：20

Chicken liver, pork liver：20

11. Sensitivity, accuracy and precision

Test Sensitivity: 1ppb

Detection limit:

Chicken, pork ……………………20ppb

Chicken liver, pork liver……………………20ppb

Accuracy:

Chicken, pork …………………………………80±10%

Chicken liver, pork liver………………………80±15%

Precision:

Variation coefficient of the ELISA kit is less than 10%.

12. Notice

12.1 The mean values of the absorbance values obtained for the standards and the samples will be reduced if the reagents and samples have not been regulated to room temperature ( 20-25℃).

12.2 Do not allow microwells to dry between steps to avoid unsuccessful reproducibility and operate the next step immediately after tap the microwells holder.

12.3. Homogenize each reagent before using.

12.4. Keep your skin away from the stop solution for it is the 2M H2SO4 solution.

12.5 Don’t use the kits out of date. Don’t exchange the reagents of different batches, or else it will drop the sensitivity.

12.6 Storage condition:

Keep the ELISA kits at 2-8℃,do not freeze. Seal rest microwell plates, Avoid straight sunlight during all incubations. Covering the microtiter plates is recommended.

12.7 Indications for the reagents going bad:

Substrate solution should be abandoned if it turns colors.

The reagents may be turn bad if the absorbance value (450/630nm) of the zero standard is less than 0.5(A450nm＜0.5).

12.8 The coloration reaction need 15min after the addition of solution A and solution B; But you can prolong the incubation time if the color is too light to be determined., never exceed 30min,On the contrary, shorten the incubation time properly.

12.9 The optimal reaction temperature is 37℃. Higher or lower temperature will lead to the changes of sensitivity and absorbance values.

13. Storage condition and storage period

Storage condition: 2-8℃.

Storage period: 6 months.