Rapid Colorimetric Determination of Acetylcholinesterase Activity

DESCRIPTION

ACETYLCHOLINESTERASE (EC 3.1.1.7, AChE), also known as RBC cholinesterase, is found primarily in the blood and neural synapses. Low serum cholinesterase activity may relate to exposure to insecticides or to one of a number of variant genotypes. AChE catalyzes the hydrolysis of the neurotransmitter acetylcholine into choline and acetic acid, a reaction necessary to allow a cholinergic neuron to return to its resting state after activation. Cholinesterase levels of cells and plasma are used as a guide in establishing safety precautions relative to exposure and contact, as well as a guide in determining the need for workers to be removed from areas of contact with the organic phosphate insecticides.

Simple, direct and automation-ready procedures for measuring AChE activity are very desirable. BioAssay Systems' QuantiChromTM Acetylcholinesterase Assay is based on an improved Ellman method, in which thiocholine produced by the action of acetylcholinesterase forms a yellow color with 5,5’-dithiobis(2-nitrobenzoic acid). The intensity of the product color, measured at 412 nm, is proportionate to the enzyme activity in the sample.

APPLICATIONS

Direct assays of acetylcholinesterase activity in blood, serum, plasma, and other biological samples. Evaluation of acetylcholinesterase inhibitors.

KEY FEATURES

Sensitive and accurate. Detection range 10 to 600 U/L AChE activity in 96-well plate assay.

Convenient. The procedure involves adding a single working reagent, and reading the optical density at 2 min and 10 min at room temperature.

High-throughput. Can be readily automated as a high-throughput 96- well plate assay for thousands of samples per day.

KIT CONTENTS (100 tests in 96-well plates)

Assay Buffer (pH 7.5): 30 mL	Reagent: 240 mg

Calibrator: 4 mL (equivalent to 200 U/L)

Storage conditions. Store all reagents at room temperature. Shelf life of at least 6 months (see expiry dates on labels).

Precautions: reagents are for research use only. Normal precautions for laboratory reagents should be exercised while using the reagents. Please refer to Material Safety Data Sheet for detailed information.

PROCEDURES

Sample preparation. Blood samples should be diluted 40-fold in the Assay Buffer, e.g. accurately pipet 5 RL blood and mix thoroughly with 195 RL Assay Buffer.

Tissue or cell lysates are prepared by brief sonication or homogenization in 0.1 M phosphate buffer (pH 7.5), followed by centrifugation at 14,000 rpm for 5 min. Use supernatant for assay.

Ideally samples should be assayed fresh. If this is not possible, refrigerate samples and assay them within 24 hours.

Reagent preparation: the Working Reagent should be prepared freshly and used within 30 min. Each reaction well requires 2 mg reagent. Calculate the amount of reagent needed and weigh this amount (mg) in a centrifuge tube. Add 200 RL Assay Buffer per 2 mg reagent. Vortex to dissolve.

1.	Calibrator: transfer 200 RL water and 200 RL calibrator separately into wells of a clear bottom 96-well plate.

Samples: add 10 RL sample per well in separate wells.

2.	Reaction: transfer 190 RL freshly prepared Working Reagent to all sample wells and tap plate briefly to mix.

Read OD412nm at 2 min and at 10 min in a plate reader.

3. Calculation: acetylcholinesterase activity is calculated as follows,

AChE Activity = O110– OD2

,.,	× n× 200 (U/L)

l/CAL - ODH2O

Where OD10 and OD2 are the OD412nm values of the sample at 10 min and 2 min, respectively. ODCAL and ODH2O are the OD412nm values of the Calibrator and water at 10 min. n is the dilution factor (n = 40 for whole blood). The number “200” is the equivalent activity of the calibrator under the assay conditions.

Note: if the calculated AChE activity is higher than 600 U/L, dilute sample in Assay Buffer and repeat this assay. Multiply the results by the dilution factor.

Unit definition: one unit of enzyme catalyzes the production of 1 Rmole of thiocholine per minute under the assay conditions (pH 7.5 and room temperature).

MATERIALS REQUIRED, BUT NOT PROVIDED

Pipeting (multi-channel) devices. Clear-bottom 96-well plates (e.g. Corning Costar) and plate reader.

GENERAL CONSIDERATIONS

1. This assay is based on an enzyme-catalyzed kinetic reaction. Addition of Working Reagent should be quick and mixing should be brief but thorough. Use of a multi-channel pipettor is recommended.

2. For assays in standard 1 mL cuvet, use 1 mL water and 1 mL Calibrator, 50 RL sample + 950 RL Working Reagent.

EXAMPLES

Two human blood samples were assayed in duplicate using the 96- well plate protocol. The AChE activities were 3,402 ± 163 and 3,660 ± 151 U/L.

4

4

Acetylcholinesterase

Acetylcholinesterase

3

3

2

2

1

1

0

0

0	10	20	30

0	10	20	30

Time (min)

Kinetics of Acetylcholinesterase Reaction in 96-well plate

LITERATURE

1.	Magnottl, RA. et al. (1987). Measurement of Acetylcholinesterase in Erythrocytes in the Field. Clin. Chem. 33/10, 1731-1 735.

2.	Kovarik, Z et al. (2003). Acetylcholinesterase active centre and gorge conformations analysed by combinatorial mutations and enantiomeric phosphonates. Biochem. J. (2003) 373, 33–40.

3.	Ordentlich, A. et al. (1996). The Architecture of Human Acetylcholinesterase Active Center Probed by Interactions with Selected Organophosphate Inhibitors. J. Biol. Chem. 271 (20): 11953–11962.